WH-082-02v02

LCD Digital Storage Oscilloscope

Manual of Operation

Applicable Models: 08201/08201B, 08202/08202B

1.

2.

Packing List

Precautions

1)

2)

Do not use a supply voltage higher than the specified maximum input voltage of

9-12VDC to the instrument, as this may damage it and also invalidate your

warranty.

Do not attempt to directly measure mains power supply without using a step down

transformer. DO NOT EXCEED 50V MAX

3.

Panel Description

Fig 1

The photo above shows the screen and various front panel controls. They are explained

respectively below.

Screen

1)

2)

3)

Vertical position indicator – the small triangle on the left border of the grid– indicates the 0 volt input

level. Its position can be changed up or down, by pressing [VPOS] and turning the adjustment knob.

Trigger level indicator – the small triangle on the right border of the grid – shows the threshold

voltage of trigger. It is triggered when the input signal level crosses the threshold voltage level; the

trigger level is referenced to the input signal ground.

Window position indicator – indicates the displayed portion of sampled memory. The distance

between the two short vertical bars at the bottom represents the whole length of sampled memory,

which is usually much longer than the screen can display. So at any given moment the screen can only

display a portion of the sampled data. The window position indicator shows which portion is being

-1-

WH-082-02v02

displayed. By pressing [HPOS] and then turning adjustment knob you can pan through the samples in

memory.

4)

5)

6)

Trigger point indicator – indicates the point where trigger occurs. Depending on the trigger position

setting the trigger point can be set to 1/8, 1/4, 1/2, 3/4, or 7/8 of the sample memory. With the trigger

point indicator users can visually identify the exact spot where trigger occurs. When the trigger point is

outside of the display window, it will be indicated by a small triangle, pointing either left or right.

“HOLD” is displayed when the oscilloscope is in HOLD state, which means capture action is halted

until HOLD state is released.

Parameter setting indicators – their functions are described below.

Fig 2

Adjustment Knob

This is rotary pot for generic setting adjustments. Turning this knob will affect the selected (highlighted)

parameter. This knob also serves as power switch. When power supply is connected pressing this knob will turn

the oscilloscope on. When the oscilloscope is on pressing and holding this knob for 2 – 3 seconds will turn the unit

off.

Coupling Select Switch

This switch selects which coupling method to be used. “DC” means input signal is directly applied to the

oscilloscope. “AC” means input signal is connected to the oscilloscope via a capacitor, which blocks the DC

component of input signal. When this switch is set to “GND” the input signal is disconnected and oscilloscope

input is connected to ground.

Vertical Sensitivity Select Switches

There are two switches for vertical sensitivity selection: The first one selects base value. The second selects

multiplier. A combination of the two settings determines actual vertical sensitivity. For example, when switch V.

SEN. 1 is in the “0.1V” position and switch V SEN 2 is in the “X2” position, the actual vertical sensitivity is (0.1 x

2 = 0.2) which means 0.2V per division.

SEC/DIV

Select time-base. When this button is pressed the time-base display will be highlighted. Time-base settings

can be adjusted by turning the adjustment knob.

V.POS

Select vertical position. When this button is pressed, the vertical position adjustment indicator will be

highlighted and vertical position can be adjusted by turning the adjustment knob.

H.POS

Select horizontal position. When this button is pressed, the horizontal position adjustment indicator will be

highlighted and the displayed portion of captured signal sample in the buffer can be adjusted horizontally by

turning the adjustment knob.

MODE

Selects trigger modes. When this button is pressed, the trigger mode adjustment indicator will be highlighted

and the trigger mode can be selected by turning the adjustment knob.

SLOPE

Selects trigger polarity. When this button is pressed, the slope select indicator will be highlighted. The

trigger slope will be toggled between rising and falling.

-2-

WH-082-02v02

LEVEL

Selects trigger level. When this button is pressed, the trigger level adjustment indicator will be highlighted

and the trigger level can be adjusted by turning the adjustment knob. Pressing this button again will toggle the

trigger source between internal and external.

HOLD

Toggle the oscilloscope between HOLD and RUN states.

4.

Connectors and Terminals

Power Supply Connector

This connector is located at bottom right of the panel. A 9 – 12V DC power supply with minimum 500mA

capacity should be used. Even though there is over-voltage protection, a power supply voltage higher than 13V

should be avoided. The centre conductor of the connector is the positive pole of the power supply.

Signal Input Connector

This connector locates at bottom left of the panel. This is a standard BNC connector for oscilloscope probe.

Please note that with a 1X probe the maximum peak-to-peak voltage applied to this input should not be greater

than 50V. If measurement for larger voltage is needed appropriate divider (use 10X probe, for example) should be

provided.

USB Connector

This connector locates at the right side of oscilloscope. This is a 5-pin mini USB B type connector. It is for

firmware upgrading and data transferring to/from PC.

M.F.T. Connector

The Multi-Function Terminal (M.F.T.) locates at the left side of oscilloscope. It serves two purposes:

1)

2)

Test Signal output

External Trigger input

The function of M.F.T. is selected by [LEVEL] key and indicated by the small letter at the right of trigger

mode indicator. The table below summarizes the functions of M.F.T.

5.

Basic Operations

Power Up

Connect power adapter to the connector at bottom-right and push the [KNOB] button. The oscilloscope first

enters bootloader and stays for 1 – 2 seconds, and waits for firmware upgrade request. If no upgrade request is

detected control is passed to main firmware. Logo and firmware version information will be briefly displayed and

then oscilloscope enters normal working state.

Set Parameters

The basic and most frequently used oscilloscope parameters are:

1)

2)

3)

Vertical related – sensitivity, vertical position [VPOS], and AC/DC coupling.

Horizontal related – Time-base [SEC/DIV] and horizontal position [HPOS].

Trigger – trigger modes [MODE], edges [SLOPE], source, and level [LEVEL].

Please refer to the descriptions in Section 4 for how to use these controls to change the parameters.

Example 1 Observe the test signal (Practice basic operations)

1)

2)

Place the tip of the probe on the centre conductor of the BNC

connecter marked M.F.T. This is the output of built-in

1KHz/5Vpp test signal. Remember to set the attenuation

switch at probe handle to 1X position.

Press the [LEVEL] button to set the trigger source to internal

(letter ‘i’ should be displayed. Otherwise you won ’t get test

signal output.)

Fig 3

-3-

WH-082-02v02

3)

4)

5)

6)

7)

8)

9)

Set vertical sensitivity to 2V (V.Sen 1 switch in position “1V” And V. Sen 2 switch in position “X2”).

Set the coupling switch to the “DC” position. (See Fig 3)

Press the [ V.POS ] button, and adjust 0V indicator to the position shown in Fig. 3.

Press the [ SEC/DIV ] button, and set the Time-base to 0.5ms. Waveform shown in Fig. 3 should be

seen. (If you find the display is unstable, you can press [LEVEL] and turn the knob so that a stable

display is obtained.)

Changing position of V. Sen 2 switch you should see that the waveform amplitude changes accordingly.

You can compare the amplitude readings.

Changing the Time-base setting to 0.2ms, for instance, you should see the signal waveform widens

accordingly. Compare readings and experiment with other settings.

Now, select “AC” with the coupling switch and you should see that the waveform is shifted down to a

position where the VPos indicator is at its midway point.

Example 2 Observe saw signal (Learn how to use trigger modes)

The schematic on the right is a simple sawtooth signal generator.

We are going to use the scope to observe its output. Build the circuit

according to the schematic. Connect power and oscilloscope as

indicated.

1)

2)

3)

4)

Select coupling switch to DC. Set V. Sen. 1 switch to the

“1V” position and V. Sen. 2 switch to the “X2” position.

(i.e. sensitivity is set to 2V/Div). Adjust the 0V to the

second last point on the vertical scale. (See Fig 4)

Power the circuit. Select Time-base to 0.5ms (or other

depending on type of transistor used). You should see a

waveform similar to that shown in Fig 4.

Select AUTO trigger mode and change the trigger level.

You should be able to observe that the waveform stabilizes

itself when it intersects with trigger level (the internal

trigger source should be selected here). Otherwise, the

waveform will jump and you won ’t be able to get a clear

view of it.

Change the trigger mode to NORM, and vary the trigger

level. Notice that the waveform only update when they

intersect with the trigger level (internal trigger source

should also be selected here). Otherwise, it will stall.

Fig 4

5)

Press the [HOLD] button to put the oscilloscope into the HOLD state. “HOLD” is displayed. Now you

can shift the waveform back and forth to view the un-displayed waveform. Pressing [HOLD] again will

release the HOLD state and put oscilloscope back into captured signal sample state.

Use 10X Probe

10X probe offers the advantage of higher input voltage range, higher

impedance, and less interference to the circuits being measured. In order to

effectively use 10X probe the compensation cap trimmer (see Fig 5) must be

calibrated. This can be done using the built-in 1KHz/5V test signal.

Set the probe attenuation to 1/10 by setting the attenuation switch in the

probe handle to the “10X” position. Insert the probe tip into the centre

conductor of M.F.T. connector. Set oscilloscope Time-base to 0.2ms/Div and

adjust trigger level so that a stable waveform display is obtained. Tune the cap

trimmer carefully (see Fig 5) so that right compensation figure is reached

(refer to Fig 6).

Fig 5

At using 10X probe remember to multiply vertical reading by 10 for true results. If, for example, the displayed

signal amplitude is 2.5V the real signal amplitude is actually 25V.

(a) Under compensation

(b) Right compensation

Fig 6

-4-

(c) Over compensation

WH-082-02v02

Vertical Position Zeroing

Sometimes you may find the vertical position indicator is not aligned to 0V level trace. This is caused by two

factors.

1)

2)

Component errors.

Interference from the ADC clock.

In order to reduce the misalignment a “Vertical Position Zeroing” function was implemented as workaround on

the DSO082. What it actually does is to add correction values to samples so that misalignment does not appear to

be that large. However, even with this correction some misalignment still exists (varies depending on Time-base

settings between 1 – 3 pixels,). This misalignment causes some inconvenience but will not affect the measurement

results. Users can use the VPos triangle as a rough indicator only. Use where the 0V level is as a reference point

and start measurements from there.

To use the “Vertical Position Zeroing” function first set the coupling switch to the “GND” position. Then press

[VPOS] button and hold it down for about 3 seconds. The VPos zeroing prompt screen will appear. Follow the

instructions on the screen to continue.

Please note that vertical position zeroing is dependent upon the individual Time-base setting. This means that

for each Time-base setting there is a specific correction value, which the zeroing process generates for that

individual Time-base. But 1ms/Div Time-base is special. Zeroing carried under 1ms/Div will set all corrections to

the same value. So for a new oscilloscope vertical position zeroing should be first done to 1ms/Div. Usually once

zeroing under 1ms/Div is done VPos alignments for most of the rest Time-base settings are ok. If the misalignment

is found too large for an individual Time-base, then you should carry out zeroing for that Time-base.

Rolling Display Mode

When Time-base is set to 0.1s/Div or slower DSO 082 uses a display mode called “Rolling”. In this mode the

trace shifts from right to left, simulating time passing.

LCD Backlight ON/OFF

Pressing [MODE] and [LEVEL] at the same time will toggle the LCD backlight ON/OFF. This is useful when

the oscilloscope is running on battery power.

Power Off

To power off the oscilloscope push down the adjustment knob and hold it for about 3 seconds. The oscilloscope

will enter stand-by mode where most of circuits are powered off. Only keypad remains active.

6.

Advanced Operations

Select Record Length

In normal running mode pushing and holding down the [HPOS] button

will make the device enter the Record Length Selection screen (see Fig 7).

Use the adjustment knob to select the record length you want and press

[HOLD] to Exit and return to normal running mode.

For most applications record length of 2000 points is enough and is thus

recommended. This will allow the oscilloscope response time to be faster. If

a longer record length is selected the slower the response will be expected

because the device needs more time load the buffer before starting a new

captured signal sample. Particularly, when the Time-base is set to 20ms/Div

or 50ms/Div and the record length of 10000 or larger is selected; you will see

a significantly slower display update.

Fig 7

Change Trigger Position

In normal running mode pushing and holding down the [LEVEL] button

will make the device enter the Trigger Position selection screen (see Fig 8).

Use the adjustment knob to select the trigger position you want and pressing

the [HOLD] button Exits and returns to normal running mode.

The Trigger position can be set to 1/8, 1/4, 1/2, 3/4, or 7/8 in the captured

signal sample buffer (regardless of record length setting). If the trigger

position is set to 1/8 this means the first 1/8 of the buffer keeps the sample

taken before trigger point, and 7/8 of the buffer keeps a sample taken after

the trigger point. If the trigger position is set to 1/2 this means buffer size                Fig 8

used to keep samples before and after trigger point is half and half. Apparently,

if you are more interested in waveforms after the trigger point you should set the trigger position to a smaller value.

And if you are more interested in a signal before the trigger then a larger trigger position setting should be used.

USB Connection

The USB connection used in 082 is a virtual Uart port based on the dedicated Uart-USB convert chip CP2102

from Silicon Lab. To use the connection, you need to install the PC drivers. Drivers for Windows and Linux can be

downloaded from the Silicon Lab’s web site.

-5-

WH-082-02v02

Download link: http://www.silabs.com/products/mcu/pages/usbtouartbridgevcpdrivers.asp                                                          x

Save captured signal samples

Captured waveforms can be stored in the on-chip EEPROM for later use.

Do This by:

1)

2)

3)

Freeze the waveform you intend to save by pressing [HOLD] key

(enter HOLD state).

Press [SLOPE] and use [KNOB] to select one of the 4 buffers.

Press [VPOS] to enter save captured signal sample screen.

4)

Press [HOLD] to save the frozen waveform to the buffer selected.

Fig 9

Recall captured signal samples

To recall saved waveforms:

1)

2)

3)

Enter HOLD state by pressing [HOLD].

Press [SLOPE] and use [KNOB] to select the buffer to be recalled.

Press [HOLD] to display the waveform in the buffer selected.

Send Screen as Bitmap File

The oscilloscope screen can be transferred to a PC as a bitmap file via the

USB. The transfer protocol used here is Xmodem. Communication format is

38400bps, 8 data bits, 1 stop bit, no parity, no flow control. To do this, please

follow the steps below:

1)

2)

3)

4)

5)

Install Virtual Com Port driver if you have not already done so.

Connect the scope to PC via a mini USB cable.

Launch Windows HyperTerminal (or any other communication tool

Fig 10

Select file extension name to be “.bmp”. Otherwise the file cannot

be opened properly. If you failed to do so in this step you can always rename the file to “.bmp” so that

it can be opened.

Put scope into HOLD state and display the waveform you are interested in on the screen.

Press the [LEVEL] button, then [HOLD]. Scope screen will be sent to PC as a bitmap file.

Using the External Trigger

1)

2)

3)

4)

To Select the trigger source, Press [LEVEL] to highlight the trigger level indicator. Press [LEVEL]

again to select the trigger source. For the external trigger, press the [LEVEL] button so that letter “e” is

displayed.

Connection of external trigger signal. Connect the external trigger signal to the M.F.T. Note: external

trigger signal amplitude should not exceed the allowed range of 0 - +10V.

Connect signal to be viewed to the oscilloscope input.

Adjust trigger level to generate trigging.

Note: External trigger signal is taken as a digital signal. The trigger threshold is LVTTL, i.e. voltage of

0.8V or lower is considered as LOW and voltage of 2.0V and higher is considered HIGH.

Using the FFT

1)

2)

3)

4)

5)

In normal oscilloscope mode. Enter FFT mode by holding the

[MODE] button down for 2 – 3 seconds.

Use [HPOS] to select FFT size (256, 512, or 1024).

Use [KNOB] to select sampling rate.

To return to oscilloscope mode press the [MODE] button again.

7.

Firmware Upgrade

Fig 11

The main firmware is run by MCU ATmega64. There are two ways to upgrade this firmware.

1)

2)

Using the pre-programmed bootloader. (See Section 5 Power up )

Using hardware programmer.

For using the bootloader you only need a miniUSB cable and an application running on a PC. This is the

easiest and saftest way so it is strongly recommended. For Windows users, the PC application is “avrubd.exe”

from****). Please download this program from JYE Tech web

site and copy it to your computer (no installation required). Refer to the document “How to Program the Scope”

for more detailed operation procedures.

-6-

WH-082-02v02

If hardware programmer is used, it is important to make sure the programmer

has a compatible pin-out. If they are different you need to re-route the signals to

match the pinouts. The programming header for ATmega64 (U2) is J3. Its pin-out

is shown Fig 12 and is also indicated on the PCB.

If JYE Tech’s USB AVR Programmer (PN: 07302) is used to program the

oscilloscope remember to connect programming ribbon cable to header J5 on

07302 board. Do not connect to J8, This is for the older 062 models only. The first

conductor of the ribbon cable (usually marked with red stripe along one edge) should be aligned to PIN 1 of

headers on both boards. Program operators should then follow the instructions of their respective programmers.

8.

Specifications

-7-

WH-082-02v02

9

Limited Warranty

JYE Tech Ltd. warrant the DSO082 oscilloscope against defects in materials and workmanship under ordinary

consumer use for a period of one year from the date of original retail purchase. During this warranty period, if a

fault occurs with the oscilloscope, and the customer follows the instructions for returning the oscilloscope, We will,

to the extent permitted by law, offer one of the following options: (i) repair the oscilloscope using either new or

refurbished parts, (ii) replace the oscilloscope with a new or refurbished one, or (iii) refund to you, all or part of the

purchase price of the oscilloscope. All replaced parts and oscilloscopes for which a refund has been given shall

become the property of JYE Tech Ltd. This limited warranty applies only to the hardware components of the

oscilloscope that have not been subjected to: accident, misuse, neglect, fire or other external causes, alternations,

repair, or commercial use.

For information on how to obtain a warranty service for your oscilloscope, please contact weihua8@weihuameter.com

10 Revision History

-8-

# Text boxes

| Items | 08201/08201B | 08202/08202B |
| --- | --- | --- |
| 082 Digital Storage Oscilloscope | 1 | 1 |
| AC/DC Power Adapter | 1 | 1 |
| BNC Probe |  | 1 |
| Operating Manual | 1 | 1 |

| Indicator | Trigger Source | Test Signal Output |
| --- | --- | --- |
| “ i “ | Internal | Enabled |
| “ e “ | External | Disabled |

that can handle the Xmodem protocol) and prepare it to receive a file.

| Vertical |
| --- |
| Number of Channels | One |
| Analog Bandwidth Range | 0 – 10MHz |
| Sensitivity | 10mV/Div – 5V/Div |
| Resolution | 8 bits |
| Input Impedance | 1M Ohm |
| Max Input Voltage | 50Vpp (for 1X probe) |
| Couple Range | DC, AC, GND |
| Horizontal |
| Max Real-time Sample Rate | 50MSPS |
| Time-base Range | 0.2us/Div – 10minute/Div |
| Record Length | 2000 – 16000 points user selectable in steps of 2000 points |
| Zoom-In | Up to 10 times in HOLD mode |
| RUN/HOLD Modes | One button switchable |
| Trigger |
| Trigger Modes | Auto, Normal, and Single |
| Trigger Polarity | Rising/Falling |
| Trigger Source | Internal, External |
| Trigger position | 1/8, 1/4, 1/2, 3/4, or 7/8 of sample buffer user adjustable |
| Trigger Point Indicator | Yes |
| External Trigger Input Range | 0 – 10V |
| External Trigger Threshold | LVTTL |
| Other Features |
| Save/recall up to 4captured signal sample s |
| Screen image hardcopy |
| Firmware upgradeable via bootloader. No programmer needed |
| USB connection for data transfer and firmware upgrade |
| FFT |
| FFT Size | 256, 512, or 1024 points selectable |
| FFT Sampling Rate | 1KSps – 50MSps adjustable |
| Window | Hamming |
| Display |
| Screen | 2-inch 128 X 64 dot-matrix black/white LCD |
| Backlight | Yes (ON/OFF controllable) |
| Contrast | Adjustable |
| Power Supply |
| Power Supply Voltage | 9 DC |
| Power Supply Current | < 200mA |
| Physical |
| Dimension | 140mm X 70mm X 30mm |
| Weight | 150 grams (not including probe and power adapter) |

Fig 12 J3 pin-out

| Version | Date | Summary |
| --- | --- | --- |
| v01 | 2011.01.23 | First release |
| v02 | 2011.01.25 | Corrected a few typo errors |