NRAC Issue #2082			December 31, 2015				12 Pages

NATIONAL RECALL ALERT CENTER - IMMEDIATE WARNING ALERT NOTIFICATION

THE FOLLOWING HAVE BEEN RECALLED OR ARE SUBJECT TO FIELD CORRECTION

THIS LISTING IS ISSUED BY - NATIONAL RECALL ALERT CENTER, WASHINGTON, DC

SUMMARY PAGE

MEDICAL DEVICES

PRODUCT		1 Step 12 Panel Cup, etc.

MANUFACTURER	Ameditech Inc.

PAGE			2

PRODUCT		Artis systems with large display

MANUFACTURER	Siemens Medical Solutions USA, Inc

PAGE			11

PRODUCT		Brilliance 64 Computed Tomography X-ray system, etc

MANUFACTURER	Philips Medical Systems (Cleveland) Inc.

PAGE			6

PRODUCT		Channel Phased Array Flex Coil, 3.0T 6

MANUFACTURER	GE Medical Systems, LLC

PAGE			10

PRODUCT		Duodenoscope

MANUFACTURER	Fujifilm Medical Systems USA, Inc

PAGE			9

PRODUCT		FB Broth, 10ml

MANUFACTURER	Hardy Diagnostics

PAGE			11

PRODUCT		Granada Biphasic Broth

MANUFACTURER	Biomerieux Inc.

PAGE			11

PRODUCT		MicroCutter Xchange 30 Blue & 30 Blue Curved Tip Cartridge

MANUFACTURER	Cardica, Inc.

PAGE			9

PRODUCT		MicroScan Pos Combo Panel, etc.

MANUFACTURER	Beckman Coulter Inc.

PAGE			6

PRODUCT		MOSAIQ Oncology Information System

MANUFACTURER	Elekta, Inc.

PAGE			8

PRODUCT		Suture Passer Black Mamba & Green Mamba

MANUFACTURER	Biomet, Inc

PAGE			10

PRODUCT		Symbiq Infusion System Pole Clamp Assembly

MANUFACTURER	Hospira, Inc

PAGE			9

PRODUCT		Syngo Imaging XS

MANUFACTURER	Siemens Medical Solutions USA, Inc

PAGE			8

PRODUCT		VAPR TRIPOLAR 90 Degree Suction Electrodes

MANUFACTURER	DePuy Mitek, Inc

PAGE			9

PRODUCT		WaterPAP Positive Airway Pressure Device

MANUFACTURER	Airways Development LLC

PAGE			8

PRODUCT		Zilver 518RX Vascular Stent

MANUFACTURER	Cook Medical

PAGE			8

PHARMACEUTICALS

PRODUCT		Daytrana

MANUFACTURER	Noven Pharmaceuticals, Inc

PAGE			11

PRODUCT		Dressol-X & Dressit-X

MANUFACTURER	Rainbow Specialty & Health Products

PAGE			11

---------------------------------------------------------------------------------------------------------------------------------

PRODUCT		1) 1 Step 12 Panel Cup

RECALL NUMBER	Z-0484-2016

PRODUCT		2) 10-Panel Pipette Drug Screen

RECALL NUMBER	Z-0485-2016

PRODUCT		3) 11 Panel Dip Card (OPI 300) w/Adult

RECALL NUMBER	Z-0486-2016

PRODUCT		4) 11 Panel Card (OPI 300) w/Adult

RECALL NUMBER	Z-0487-2016

PRODUCT		5) 12-Panel Dip Drug Screen with adulterant

RECALL NUMBER	Z-0488-2016

PRODUCT		6) 6 Panel Cup Body and Lid w/Adult (AU)

RECALL NUMBER	Z-0489-2016

PRODUCT		7) BZO Strip Dip Card (300ng/ml)

RECALL NUMBER	Z-0490-2016

PRODUCT		8) CLIA RDTC 12-Panel Cup (w/Adulteration)

RECALL NUMBER	Z-0491-2016

PRODUCT		9) DrugCheck Waived 12 Panel Cup w/Adult

RECALL NUMBER	Z-0492-2016

PRODUCT		10) DrugSmart 10 Test Cup (PCP & OXY)

RECALL NUMBER	Z-0493-2016

PRODUCT		11) DrugSmart 11 Test Cup

RECALL NUMBER	Z-0494-2016

PRODUCT	12) DrugSmart 12 Test Cup (Multi & BUP) & DrugSmart 12 Test Cup (Multi/Single/BUP)

RECALL NUMBER	Z-0495-2016

PRODUCT		13) DrugSmart 9 Test Cup (BUP)

RECALL NUMBER	Z-0496-2016

PRODUCT		14) DrugSmart Cup 10

RECALL NUMBER	Z-0497-2016

PRODUCT		15) DrugSmart Cup THC/COC/AMP//MET/OPI2/BZO

RECALL NUMBER	Z-0498-2016

PRODUCT		16) DrugSmart Dip (MET/OPI2)+AMP+BZO+COC+THC

RECALL NUMBER	Z-0499-2016

PRODUCT		17) DrugSmart Dip AMP/OPI2/COC/THC/BZO Test

RECALL NUMBER	Z-0500-2016

PRODUCT		18) EDI BZO Uncut Sheet - Cup

RECALL NUMBER	Z-0501-2016

PRODUCT		19) iCup DX 10 Panel w/Adulterant

RECALL NUMBER	Z-0502-2016

PRODUCT		20) iCup DX 11 Panel w/ Adulterant

RECALL NUMBER	Z-0503-2016

PRODUCT		21) iCup DX 12 Panel w/ Adulterant

RECALL NUMBER	Z-0504-2016

PRODUCT		22) iCup DX 6 Panel w/Adulterant

RECALL NUMBER	Z-0505-2016

PRODUCT		23) iCup DX 8 Panel w/Adulterant

RECALL NUMBER	Z-0506-2016

PRODUCT		24) Immutest 10 Pnl Drug Screen Cup w/ Adult

RECALL NUMBER	Z-0507-2016

PRODUCT		25) Immutest 11 Pnl Drug Screen Cup w/ Adult

RECALL NUMBER	Z-0508-2016

PRODUCT		26) Immutest 5 Pnl Drug Cup w/ Adult -waived

RECALL NUMBER	Z-0509-2016

PRODUCT		27) Immutest 12 Panel Cup w/Adulteration

RECALL NUMBER	Z-0510-2016

PRODUCT		28) ImmuTest 12 Panel Dip Card

RECALL NUMBER	Z-0511-2016

PRODUCT		29) ImmuTest Cup 7-panel w/ Adulteration

RECALL NUMBER	Z-0512-2016

PRODUCT		30) Immutest Cup CLIA Waived w/Adulteration

RECALL NUMBER	Z-0513-2016

PRODUCT		31) Micro Screen 6 Panel Cup with Adult (AU)

RECALL NUMBER	Z-0514-2016

PRODUCT		32) ProScreen 10 Drug Panel Cassette Test

RECALL NUMBER	Z-0515-2016

PRODUCT		33) ProScreen 10 Panel Cup

RECALL NUMBER	Z-0516-2016

PRODUCT		34) ProScProScreen 10 Panel Cup with Adulterants

RECALL NUMBER	Z-0517-2016

PRODUCT		35) ProScreen 10 Panel Dip Card

RECALL NUMBER	Z-0518-2016

PRODUCT		36) ProScreen 10 Panel Dip Card w/Adult

RECALL NUMBER	Z-0519-2016

PRODUCT		37) ProScreen 12 Drug Cup w/Adulteration

RECALL NUMBER	Z-0520-2016

PRODUCT		38) ProScreen 12 Panel Cup w/Adult

RECALL NUMBER	Z-0521-2016

PRODUCT		39) ProScreen 12 Panel Dip Card

RECALL NUMBER	Z-0522-2016

PRODUCT		40) ProScreen 5 Panel Dip Card

RECALL NUMBER	Z-0523-2016

PRODUCT		41) ProScreen 5 Panel Dip Card w/Adult

RECALL NUMBER	Z-0524-2016

PRODUCT		42) ProScreen 6 Panel Cassette w/Adult  (AU)

RECALL NUMBER	Z-0525-2016

PRODUCT		43) ProScreen 6 Panel Cup w/Adult (AU)

RECALL NUMBER	Z-0526-2016

PRODUCT		44) ProScreen 6 Panel Dip Card

RECALL NUMBER	Z-0527-2016

PRODUCT		45) ProScreen 6 Panel Dip Card w/Adult  (AU)

RECALL NUMBER	Z-0528-2016

PRODUCT		46) ProScreen 6 Panel Dip Drug Screen

RECALL NUMBER	Z-0529-2016

PRODUCT		47) ProScreen 7 Panel Dip Card

RECALL NUMBER	Z-0530-2016

PRODUCT		48) ProScreen 8 Panel Dip Card

RECALL NUMBER	Z-0531-2016

PRODUCT		49) ProScreen 8 Panel Dip Card w/Adult

RECALL NUMBER	Z-0532-2016

PRODUCT		50) ProScreen 8 Panel Dip Card w/Adult

RECALL NUMBER	Z-0533-2016

PRODUCT		51) ProScreen 9 Panel Cup w/Adulterants

RECALL NUMBER	Z-0534-2016

PRODUCT		52) ProScreen CLIA Waived Cup 6 Drugs w/Adult

RECALL NUMBER	Z-0535-2016

PRODUCT		53) ProScreen CLIA Waived Cup  with 6 Drugs

RECALL NUMBER	Z-0536-2016

PRODUCT		54) ProScreen CLIA Waived Cup  with 8 Drug

RECALL NUMBER	Z-0537-2016

PRODUCT		55) ProScreen CLIA Wvd 12 Panel Cup w/Adult

RECALL NUMBER	Z-0538-2016

PRODUCT		56) ProScreen Cup CLIA with 6 Panel Drug

RECALL NUMBER	Z-0539-2016

PRODUCT		57) RealityCHECK Tilt Cup 12 Panel + Adult

RECALL NUMBER	Z-0540-2016

PRODUCT		58) THERMO FISHER 6 DRUG CASSETTE

RECALL NUMBER	Z-0541-2016

PRODUCT		59) THERMO FISHER 8 DRUG CASSETTE

RECALL NUMBER	Z-0542-2016

CODE	1) Item No. NBCA-12M-W, Product code: 143107, 151657, 153155, 153529; 2) Item No. PSP-10M, product code: 153135, 154119; 3) Item No. 11125AA, product code: 153632; 4) Item No. 71125AA, product code: 144376; 5) Item No. PSDA-12BUP, PSD-6MTDBO300, PSD-10PPX, PSD-10MOX, 60960D, 61127D, DCC-81205-5, IMCA-10M5, product code: 151688, 152423, 153307, 153912, 153921, 153961, 153962, 154104, 154112, 154236, 154264; 6) Item No. TDDA-6MBAU-CUP, product code: G140255-C, G140422-C, G150063-C, G150184-C, G150211-C, G150226-C, G150240-C, G150286-C; 7) Item No. 100170, product code: 141233, 141669, 141790, 141820, 141958, 142082, 142531, 142532, 142835, 143020, 143471, 143553, 152021, 152432, 152546, 153094, 153538, 152104; 8) Item No. CLIA-RDTC-12, DCC-81205-5, product code: 142709, 142840, 143008, 143571, 143665, 144217, 152234, 152379; 9) Item No. DCC-81205-5, product code: 143235, 144216, 151591, 153132, 153489; 10) Item No. 61020D, product code: 143978, 144485, 151982, 153095; 11) Item No. 61127D, product code: 144196, 144375, 150096, 151980, 152737; 12) Item No. 61210D, 61205D, product code: 152736, 143864, 150204, 151981; 13) Item No. 60960D, product code: 144029, 152738; 14) Item No. 61093D-3C, 61085D, 61015D, product code: 143516, 143977, 150397, 152070, 152751; 15) Item No. 60600D, product code: 150309; 16) Item No. 30600D, product code: 142122, 142699, 143170; 17) Item No. 30520D, product code: 142141, 143169; 18) Item No. 160170, product code: 141260, 141294, 141909, 141925, 142074, 142243, 142428, 142614, 142712, 142911, 144555, 151925, 152003, 152223, 152498, 152723, 152897, 153311, 153559, 153818, 154184, 152103; 19) Item No. I-DXA-1107-142, product code: 143037, 143106, 144152, 144218, 152456, 152457, 152698, 153048, 153129, 153201, 153385, 153648; 20) Item No. I-DXA-1117-131, product code: 152459, 152823; 21) Item No. I-DXA-1127-023, product code: 143345, 143442, 151537, 152163, 152334, 152674, 152697, 153075, 153127, 153216, 153390, 153499, 153634; 22) Item No. I-DXA-167-013, product code: 141900, 142194, 142959, 144377, 151998, 152740, 152993, 153229, 153363, 153544; 23) Item No. I-DXA-187-016, product code: 141971, 142195, 142397, 142674, 142997, 144209, 144463, 150018, 150883, 151525, 151526, 152180, 152741, 153045, 153150, 153151, 153211, 153376, 153650; 24) Item No. IMCA-10M5, product code: 143846, 144380, 150312, 150505, 153177, 153488; 25) Item No. IMCA-11OP, product code: 144103; 26) Item No. IMCA-5AB-W, product code: 141758, 141921, 142657, 143178, 143581, 144563, 151337, 152240, 152782, 153154, 153401; 27) Item No. IMCA-11OP, product code: 143847, 144565, 150313, 152169, 152338; 28) Item No. IMD-12MO3, product code: 144009, 144078, 144099, 144125, 144193, 144268, 152996, 153077; 29) Item No. IMCA-7P, product code: 143861, 144204, 144263; 30) Item No. IMCA-7MT-W, product code: 143937, 144574, 145061; 31) Item No. MSCA-6MBAU, product code: G140256, G140281, G140307, G140340, G140418, G140419, G150135, G150141, G150187, G150257; 32) Item No. PSP-10PPX, product code: 152164, 153134; 33) Item No. PSCup-10M, product code: 142596, 142984, 142985, 143144, 143145, 143146, 143240, 143241, 143351, 143875, 144102, 144106, 144239, 144241, 150651, 151706, 152165, 152167, 152168, 153102, 153107, 153251, 153411, 153412, 153633; 34) Item No. PSCupA-10AM, product code: 144240, 150059; 35) Item No. PSD-10MOX, PSD-10MOB, PSD-10MMO300, product code: 141963, 142196, 143166, 143475, 143827, 152038, 152172, 152218, 152323, 152393, 152394, 152870, 153362; 36) Item No. PSDA-10MO, PSDA-10MMO300, PSDA-10PPX, PSDA-10BUP, PSD-12BUP, PSD-10M, PSCupA-6MB-W, PSCup-10MMO300, PSD-5AB, PSD-10CMO3, PSCupA-10MO, PSCupA-12M-W, I-DXA-1127-023, PSCupA-12TBU, product code: 144902, 150474, 151686, 151687, 152043, 152225, 152312, 152402, 152952, 153306, 153563, 153663, 153664, 153665, 153680, 153682, 153703, 153704, 153759, 153760, 153763, 153852, 153853; 37) Item No. PSCupA-12TBU, product code: 142986, 143352, 151745, 152170, 153262; 38) Item No. PSCupA-12M, product code: 142570, 144104; 39) Item No. PSD-12BUP, product code: 152401; 40) Item No. PSD-12BUP, product code: 141953, 151770, 151771, 153522; 41) Item No. PSDA-5MB300, product code: 152912; 42) Item No. PSPA-6MBAU, product code: G150179; 43) Item No. PSCupA-6MBAU, product code: G140258, G140313, G140314, G140368, G140384, G140389, G140399, G140426, G150185, G150199, G150246, G150307; 44) Item No. PSCuPSD-6BUP300, PSD-6MTDBO300, PSD-6BUPO, PSD-6MBO, PSD-6MB, product code: 141955, 142043, 142055, 142161, 142513, 142600, 142824, 142937, 142938, 142969, 143143, 143149, 143234, 143513, 143758, 143806, 143807, 144441, 152035, 152042, 152075, 152911, 153237, 153240, 153523; 45) Item No. PSDA-6MBAU, product code: G150001, G150244, G150293; 46) Item No. PSD-6BUPO3, product code: 152112, 152705, 153400; 47) Item No. PSD-7MPB, PSD-7MO, PSD-7M, product code: 142059, 143006, 143570, 152175, 153524; 48) Item No. PSD-8P, product code: 152305, 152306, 153503; 49) Item No. PSDA-8P, PSDA-8MBO300, product code: 152950, 152951, 153325; 50) Item No. PSDA-8P, PSDA-8MBO300, product code: 152950, 152951, 153325; 51) Item No. PSCUPA-9BO300, product code: G150252; 52) Item No. PSCupA-6MB-W, product code: 141906, 143153, 143803, 150003, 150207, 150650, 151786, 152081; 53) Item No. PSCup-6BO-W, product code: 141779, 142656, 142963, 143874, 144562, 151996; 54) Item No. PSCup-8P-W, product code: 142964, 150790, 151707; 55) Item No. PSCupA-12M-W, product code: 152171, 152216, 153103; 56) Item No. PSCup-6MB-W, product code: 142153, 142594, 143825, 144227, 152005, 152972; 57) Item No. RC12TCA, product code: 153709; 58) Item No. MTPA-6MBAU, product code: G150189, G150271, G150299; 59) Item No. MTPA-8MBAU, product code: G150328

MANUFACTURED & RECALLED BY	Ameditech Inc, San Diego, CA

QUANTITY	1) 1,377 kits, 2) 240 kits, 3) 40 kits, 4) 50 kits, 5) 858 kits, 6) 3,603 kits, 7) 31,472 kits, 8) 10,320 kits, 9) 1,907 kits, 10) 340 kits, 11) 320 kits, 12) 160 kits, 13) 320 kits, 14) 272 units, 15) 60 kits, 16) 120 kits, 17) 80 units, 18) 45,584 kits, 19) 6,304 kits, 20) 26 units, 21) 9,566 kits, 22) 2,085 kits, 23) 8,290 units, 24) 320 kits, 25) 240 kits, 26) 1,909 kits, 27) 200 kits, 28) 2,693 kits, 29) 60 kits, 30) 144 kits, 31) 4,646 kits, 32) 160 kits, 33) 3,411 kits, 34) 96 kits, 35) 2,452 kits, 36) 1,989 kits, 37) 636 kits, 38) 1200 kits, 39) 60 kits, 40) 338 kits, 41) 120 kits, 42) 60 kits, 43) 7,805 kits, 44) 2,675 kits, 45) 168 kits, 46) 200 kits, 47) 567 kits, 48) 249 kits, 49) 82 kits, 50) N/A, 51) 3 kits, 52) 2,136 kits, 53) 492 kits, 54) 129 kits, 55) 1,680 kits, 56) 743 kits, 57) 80 kits, 58) 300 kits, 59) 100 kits

DISTRIBUTION		Nationwide

REASON	Ameditech’s Drugs of Abuse Tests listed above have shown reduced reactivity for up to two (2) of the nineteen (19) Benzodiazepine (BZO) compounds for which specificity information is provided in the product insert. The particular BZO compounds for which there may be reduced reactivity are Clonazepam and Chlordiazepoxide.

---------------------------------------------------------------------------------------------------------------------------------

PRODUCT		1) MicroScan Pos Combo Panel Type 21

RECALL NUMBER	Z-0558-2016

PRODUCT		2) MicroScan Pos Combo Panel Type 33

RECALL NUMBER	Z-0559-2016

PRODUCT		3) MicroScan Pos Combo Panel Type 34

RECALL NUMBER	Z-0560-2016

PRODUCT		4) MicroScan Pos Breakpoint Combo Panel Type 20

RECALL NUMBER	Z-0561-2016

PRODUCT		5) MicroScan Pos Breakpoint Combo Panel Type 23

RECALL NUMBER	Z-0562-2016

CODE	1) Part number B1017-201, lot 2016-05-05; 2) Part number B1017-211, lot 2016-04-29; 3) Part number B1017-214, lot 2016-04-29, 2016-05-06; 4) Part number B1017-202, lot 2016-04-30; 5) Part number B1017-206, lot 2016-05-06

MANUFACTURED & RECALLED BY	Beckman Coulter Inc., Brea, CA

QUANTITY		1) 96 units, 2) 1,276 units, 3) 1,215 units, 4) 701 units, 5) 616 units

DISTRIBUTION		Nationwide

REASON	Beckman Coulter has confirmed falsely negative results for Voges-Proskauer (VP) (pale-pink/brown/colorless) with Quality Control (QC) American Type Culture Collection (ATCC) organism Staphylococcus aureus ATCC 29213 in a portion of affected lots of MicroScan Pos Combo and MicroScan Pos Breakpoint Combo panels. The expected result is positive.

---------------------------------------------------------------------------------------------------------------------------------

PRODUCT		1) Brilliance 64 Computed Tomography X-ray system

RECALL NUMBER	Z-0544-2016

PRODUCT		2) Brilliance iCT Computed Tomography X-ray system

RECALL NUMBER	Z-0545-2016

PRODUCT		3) Brilliance iCT SP  Computed Tomography X-ray system

RECALL NUMBER	Z-0546-2016

PRODUCT		4) Ingenuity Core Computed Tomography X-ray system

RECALL NUMBER	Z-0547-2016

PRODUCT		5) Ingenuity Core 128 Computed Tomography X-ray system

RECALL NUMBER	Z-0548-2016

PRODUCT		6) Ingenuity CT Computed Tomography X-ray system

RECALL NUMBER	Z-0549-2016

CODE			1) Model No. 728231; S/N: 4001, 9074, 9099, 9538, 9549, 9573, 9638, 9646, 9656, 9667, 9703, 9738, 9755, 9782, 9811, 9822, 9909, 9923, 9927, 9933, 10021, 10042, 10077, 10150, 10180, 10379, 29056, 29107, 29115, 90124, 90154, 95024, 95130, 95222, 95236, 95258, 95261, 95364, 95399, 95410, 95475, 95551, 95614, 95668, 95688, 95691; 2) Model No. 728306; S/N: 85015, 85021, 85028, 85034, 85036, 85039, 85042, 85047, 85051, 100030, 100038, 100040, 100051, 100054, 100058, 100065, 100072, 100093, 100094, 100099, 100101, 100140, 100144, 100162, 100181, 100190, 100214, 100217, 100228, 100236, 100244, 100278, 100421, 100462, 100472, 100613, 100616, 100624, 100627, 100628, 100632, 100639, 100640, 100641, 100644, 100647, 100651, 100654, 100655, 100659, 100660, 100662, 100667, 100669; 3) Model No. 728311; S/N: 200047, 200051, 200067, 200074, 200098, 200127, 200200, 200202, 200203, 200204, 200206, 200209; 4) Model No. 728321; S/N: 52007, 52008, 52010, 52013, 52014, 52015, 52018, 52019, 52020, 52021, 52024, 52027, 52032, 52033, 52034, 52035, 52036, 52037, 52038, 52040, 52041, 52042, 52045, 52046, 52047, 52048, 52051, 52053, 52055, 52057, 52058, 52060, 52064, 52065, 52067, 52068, 52069, 52070, 52075, 52078, 310017, 310026, 310039, 310042, 310052, 310055, 310059, 310064, 310069, 310071, 310081, 310088, 310098, 310111, 310113, 310119, 310121, 310133, 310143, 310144, 310149, 310151, 310157, 310159, 310160, 310204, 310205, 310206, 310208, 310209, 310210, 310211, 310212, 310213, 310215, 310216, 310217, 310219, 310220, 310223, 310224, 310225, 310226, 310228, 310231, 333002, 333009, 333014, 333015, 333030, 333031, 333037, 333042, 333043, 333044, 333046, 333048, 333049, 333050, 333052, 333053, 333054, 333055, 333056, 333058, 333059, 333060, 333061, 333062, 333063, 333065, 333066, 333067, 333069, 333070, 333072, 333073, 333074, 333075, 333076, 333077, 333078, 333079, 333080, 333081, 333082, 333086, 333087, 333090, 333091, 333093, 333101, 333110, 333112; 5) Model No. 728323; S/N: 30016, 32012, 32013, 32016, 32024, 32028, 32033, 32036, 32041, 32046, 32048, 32050, 32051, 32052, 32053, 32054, 32055, 32056, 32058, 32059, 32062, 32063, 32064, 32068, 32069, 32070, 32072, 32073, 32074, 32075, 32077, 32078, 32079, 32081, 32082, 32083, 32084, 32086, 32087, 32089, 32092, 32093, 32095, 32097, 32098, 32099, 32100, 32102, 32103, 32104, 32105, 32106, 32108, 32110, 32113, 32114, 32115, 32116, 32119, 32120, 32122, 32123, 32124, 32125, 32128, 32129, 32131, 32132, 32133, 32134, 32135, 32136, 32137, 32138, 32140, 32141, 32145, 32146, 320003, 320006, 320018, 320024, 320029, 320030, 320033, 320037, 320054, 320058, 320069, 320073, 320078, 320079, 320090, 320092, 320099, 320114, 320115, 320123, 320125, 320126, 320128, 320133, 320134, 320138, 320146, 320200, 320201, 320202, 320203, 320204, 320205, 320208, 320212, 320217, 320219, 320220, 320222, 320223, 320224, 320225, 320226, 320228, 320231, 320232, 320239, 320241, 320243, 320246, 320247, 336013, 336015, 336016, 336018, 336027, 336029, 336034, 336046, 336059, 336060, 336064, 336066, 336068, 336071, 336072, 336073, 336075, 336078, 336079, 336080, 336081, 336083, 336084, 336085, 336086, 336087, 336088, 336089, 336091, 336092, 336095, 336097, 336098, 336099, 336105, 336107, 336108, 336110, 336111, 336112, 336113, 336115, 336116, 336117, 336119, 336120, 336121, 336122, 336123, 336126, 336128, 336131, 336132, 336133, 336135, 336137, 336138, 336139, 336140, 336143, 336144, 336145, 336147, 336148, 336150,  336151; 6) Model No. 728326; S/N: 2, 1234, 30005, 30009, 30017, 30018, 30020, 30023, 30028, 30031, 30034, 30037, 31001, 32065, 32067, 52030, 52031, 52044, 52049, 52050, 300159, 300204, 300205, 300207, 300208, 300209, 310002, 310003, 333057, 333064, 336061

MANUFACTURED & RECALLED BY	Philips Medical Systems (Cleveland) Inc, Cleveland, OH

QUANTITY		1) 30 units, 2) 51 Units, 3) 12 Units, 4) 135 Units, 5) 177 Units, 6) 31 Units

DISTRIBUTION		Nationwide and internationally

REASON	Software versions preceding software version 4.1.2 provided a basic feature of axial perfusion scan. Software version 4.1.2 combined the basic feature of axial perfusion scan with an advanced optional feature of helical perfusion scan made available via a license key resulting in the basic feature not being available by default.

---------------------------------------------------------------------------------------------------------------------------------

PRODUCT		WaterPAP Positive Airway Pressure Device

RECALL NUMBER	Z-0552-2016

CODE			Reference #WP-7700; Lot 8605A

MANUFACTURED & RECALLED BY	Airways Development LLC, Kenilworth, NJ

QUANTITY		1,464 units

DISTRIBUTION		Nationwide

REASON	Airways Development LLC has received a complaint about a canister leaking from WaterPAP Lot 8605A.  During a review with their supplier it could not be determined if this was an isolated incident or if the whole lot is impacted.

---------------------------------------------------------------------------------------------------------------------------------

PRODUCT		MOSAIQ Oncology Information System

RECALL NUMBER	Z-0557-2016

CODE			Versions 2.50 and higher

MANUFACTURED/RECALLED BY	Elekta, Inc., Atlanta, GA

QUANTITY		399

DISTRIBUTION		Nationwide and internationally

REASON	Incorrect drug dosage due to "Age Limit" and patient weight data item issue.

---------------------------------------------------------------------------------------------------------------------------------

PRODUCT		Syngo Imaging XS is a Picture Archiving and Communication System (PACS)

RECALL NUMBER	Z-0550-2016

CODE			Model number 10558586, Software version VB10C

MANUFACTURED & RECALLED BY	Siemens Medical Solutions USA, Inc, Malvern, PA

QUANTITY		42

DISTRIBUTION		Nationwide

REASON	For the Syngo Workflow SLR System with Software Ver: VB10C: Printouts may be printed in incorrect anatomical size when using syngo Imaging XS filming application in conjunction with a printer not released for anatomical print usage. This may happen due to an invalid combination of printer and the syngo Imaging XS printer configuration file.

---------------------------------------------------------------------------------------------------------------------------------

PRODUCT		Zilver 518RX Vascular Stent with Rapid Exchange Delivery System

RECALL NUMBER	To be determined

CODE			Lot numbers:

CF997821

CF997822

CF997823

CF997824

CF997825

CF997826

CF997827

CF997828

CF997829

CF997830

CF997831

CF997832

CF997833

CF997834

CF997835

CF997836

CF997837

CF997838

CF997839

CF997840

CF997841

CF997842

CF997843

CF997844

CF997845

CF997846

CF997847

CF997848

CF997849

CF997850

CF997851

CF997852

CF997853

CF997854

CF997855

CF997856

CF997857

CF997858

CF997859

CF997860

CF997861

CF997862

CF997863

MANUFACTURED/RECALLED BY	Cook Medical, Bloomington, IN

QUANTITY	Not available at publication

DISTRIBUTION	Nationwide

REASON		Due to potential non-conformance of the delivery system.

-----------------------------------------------------------------------------------------------------------------------------------------------

-----------------------------------------------------------------------------------------------------------------------------------------------

PRODUCT		Symbiq™ Infusion System Pole Clamp Assembly

RECALL NUMBER	Field correction

CODE	List numbers 16026-XX-XX and 16027-XX-XX, all serial numbers; pole clamp assembly part number HSP3313-E12, lot number 481976

MANUFACTURED & RECALLED BY	Hospira, Inc, Lake Forest, IL

QUANTITY		Not available at publication

DISTRIBUTION		Nationwide

REASON		The potential for Symbiq pole clamp assemblies to be missing the rubber stop pad.

-----------------------------------------------------------------------------------------------------------------------------------------------

PRODUCT		Duodenoscope

RECALL NUMBER	Safety notice

CODE			Model number ED-530XT

MANUFACTURED & RECALLED BY	Fujifilm Medical Systems USA, Inc, Wayne, NJ

QUANTITY		Not available at publication

DISTRIBUTION		Nationwide

REASON	Fujifilm Medical Systems USA issued revised, validated manual reprocessing instructions for the ED-530XT duodenoscope to replace those provided in the original device labeling. The revised instructions include a more rigorous protocol for pre-cleaning, manual cleaning and high-level disinfection procedures.

-----------------------------------------------------------------------------------------------------------------------------------------------

PRODUCT		1) MicroCutter Xchange 30 Blue Cartridge

RECALL NUMBER	Z-0482-2016

PRODUCT		2) MicroCutter Xchange 30 Blue Curved Tip Cartridge

RECALL NUMBER	Z-0483-2016

CODE			1) Model number: FG-025320, Lot numbers:

120416D

120521E

120606F

120613B

120621A

120910A

120926H

121005C

121022A

121031A

121119B

121217F

130131A

130207A

130213A

130220H

130227C

130307B

130320A

130328A

130426A

130508A

130516A

130619B

130613H

130716C

130815E

130822C

130911C

131003E

131113C

131125D

131212F

140109F

140115D

140127B

140204B

140226J

140304A

140310K

140317B

140325B

140430C

140507B

140604C

140610E

140617H

140624E

140707G

140710D

140715H

140804D

140819B

140828C

140916E

140922C

140926C

141008E

141016E

141103A

141106A

150415D

150625B

150813B

151006G

2) Model Number: FG-025321, Lot numbers:

150309A

150527D

150603F

150617B

150820A

150902A

150909B

150924A

MANUFACTURED & RECALLED BY	Cardica, Inc., Redwood City, CA

QUANTITY		568 boxes or 6816 cartridges

DISTRIBUTION		AL, CA, FL, GA, IA, MA, MD, MN, NY, OH, PA, RI, TX, WA, WI

REASON	Cardica, Inc. has received reports that use of the Blue Cartridges with the MicroCutter XCHANGE 30 Stapler was associated with incomplete firing of staples, potentially resulting in an incomplete transection or anastomosis of tissue during the surgical procedure.

---------------------------------------------------------------------------------------------------------------------------------

PRODUCT		VAPR® TRIPOLAR 90" Degree Suction Electrodes

RECALL NUMBER	Z-0479-2016

CODE			Model/Catalog No 225028; U1509202, U1509162

MANUFACTURED & RECALLED BY	DePuy Mitek, Inc, a Johnson & Johnson Co, Raynham, MA

QUANTITY		127

DISTRIBUTION		AL, CA, CT, FL, KY, MN, MS, ND, NJ, NM, TX

REASON	Mitek identified that on VAPR Tripolar 90 Degree Suction Electrode the ablation and coagulation buttons are colored incorrectly.

---------------------------------------------------------------------------------------------------------------------------------

PRODUCT		1) Suture Passer, Black Mamba

RECALL NUMBER	Z-0471-2016

PRODUCT		2) Suture Passer, Green Mamba

RECALL NUMBER	Z-0472-2016

CODE	1) Catalog Number 110010849; Lot Number Identification 167880, 169620, 253190, 520810, 538490, 551550, 716270, 960180; 2) Catalog Number 110010850; Lot Number Identification 167890, 231120, 253210, 550050, 551580, 716300, 953230

MANUFACTURED & RECALLED BY	Biomet, Inc, Warsaw, IN

QUANTITY		1) 50, 2) 13

DISTRIBUTION		AL, AR, CA, FL, GA, HI, IN, KS, KY, LA, MN, MO, NC, NJ, NY, OH, TX, UT, VA, WI

REASON	There may be a separation and an insufficient weld between the spring and washer. The weld between the spring and washer may fracture, causing the Mamba Suture Passer instrument’s jaw to not actuate properly.

---------------------------------------------------------------------------------------------------------------------------------

PRODUCT		Channel Phased Array Flex Coil, 3.0T 6

RECALL NUMBER	Z-0553-2016

CODE			Product number: M0050SS; Lot number:

00000001140049

00000001140050

0000011-4-0050

0000011-4-0009

0000011-4-0015

0000011-4-0017

0000011-4-0018

0000011-4-0024

0000014-4-0007

0000011-4-0028

0000011-4-0032

0000011-4-0046

0000011-4-0047

0000028-3-0011

0000044-4-0046

0000011-4-0064

0000011-4-0065

0000011-4-0066

0000011-4-0067

0000011-4-0075

0000011-4-0076

011-4-0085

011-4-0170

0000011-4-0094

011-4-0118

0000011-4-0101

0000011-4-0102

0000011-4-0107

0000011-4-0108

0000011-4-0114

0000011-4-0115

011-4-0117

011-4-0172

011-4-0119

011-4-0120

011-4-0124

077-4-0123

011-4-0125

011-4-0126

011-4-0129

011-4-0130

011-4-0146

011-4-0149

011-4-0154

011-4-0155

011-4-0162

011-4-0164

0000011-4-0055

0000011-4-0062

0000011-4-0080

0000011-4-0106

011-4-0169

0000011-4-0019

0000011-4-0020

011-4-0128

11-4-0127

0000011-4-0027

0000011-4-0033

0000011-4-0133

0000011-4-0134

0000011-4-0086

0000011-4-0092

0000011-4-0153

0000011-4-0158

0000011-4-0159

0000011-4-0103

0000011-4-0104

0000011-4-0156

0000011-4-0157

0000011-4-0165

0000011-4-0166

0000028-4-0016

0000011-4-0060

0000011-4-0061

0000011-4-0093

0000011-4-0098

0000011-4-0034

0000011-4-0035

0000011-4-0056

0000011-4-0057

0000011-4-0030

0000011-4-0031

0000011-4-0042

0000011-4-0043

0000011-4-0044

0000011-4-0045

0000011-4-0143

0000011-4-0144

MISSING_SN_GON4217108

Missing_SN_GON4227287

MANUFACTURED/RECALLED BY	GE Medical Systems, LLC, Waukesha, WI

QUANTITY		54

DISTRIBUTION		AZ, CA, CO, IL IN, MI, MN, NY, OH, OK, SD, TX, WI

REASON	Coil overheating can occur when the device is used in Mode 2 setup. This could lead to a serious patient thermal injury.

---------------------------------------------------------------------------------------------------------------------------------

PRODUCT		Granada Biphasic Broth

RECALL NUMBER	Z-0556-2016

CODE			REF 42722, Lot Numbers 896, 897, 899

MANUFACTURED/RECALLED BY	Biomerieux Inc, Hazelwood, MI

QUANTITY		24 cartons

DISTRIBUTION		CA, GA, IN, MA, MN, MO, NY, OH, PA

REASON	FDA has determined that the product requires 510(k) clearance to continue marketing/distribution in the United States.

-----------------------------------------------------------------------------------------------------------------------------------------------

PRODUCT		Artis systems with large display

RECALL NUMBER	Z-0481-2016

CODE	Model# 10094135, 10094137, 10094141, 10848280, 10848281, 10848282, 10848355

MANUFACTURED & RECALLED BY	Siemens Medical Solutions USA, Inc, Malvern, PA

QUANTITY		17

DISTRIBUTION		AL, CA, FL, IL, MA, MI, ND, NJ, OH, TN, TX, WI

REASON	Due to improper soldering in a limited number of large display bypass module components of a specific production lot, a loss of video signal could potentially occur. This has not been observed in the field and only sporadic cases have been observed in factory screening.

-----------------------------------------------------------------------------------------------------------------------------------------------

PRODUCT		FB Broth, 10ml

RECALL NUMBER	Z-0470-2016

CODE			Catalog number K31, lot number 15231

MANUFACTURED & RECALLED BY	Hardy Diagnostics, Santa Maria, CA

QUANTITY		500 25pk/20

DISTRIBUTION		CA, MD, UT, WI

REASON	Due to potential Burkholderia fungorum contamination.

-----------------------------------------------------------------------------------------------------------------------------------------------

PRODUCT		1) Dressol-X

RECALL NUMBER	Z-0476-2016

PRODUCT		2) Dressit-X

RECALL NUMBER	Z-0477-2016

CODE	1) Radiopaque D-502; NDC 11004-502-40; Lots 14317, 15007, 15140; 2) Radiopaque D-545; NDC 11004-545-40; lot 5012

MANUFACTURED & RECALLED BY	Rainbow Specialty & Health Products, Markham, Ontario, Canada

QUANTITY		1) 2775, 2) 150

DISTRIBUTION		Nationwide

REASON		The labeling does not currently list the percentage of drug components.

-----------------------------------------------------------------------------------------------------------------------------------------------

PRODUCT		Daytrana® (methylphenidate transdermal system)

RECALL NUMBER	To be determined

CODE	NDC 68968-5555-3, 68968-5554-3, 68968-5553-3; lot numbers 77379, 77140, 77380, 77303, 76466, 76444, 76486, 76443, 76701, 76674, 76911, 76912, 76910, 77161, 77378, 77305, 77546, 78068, 77396, 77879, 77871, 77878, 78037, 76913, 76944

MANUFACTURED & RECALLED BY	Noven Pharmaceuticals, Inc, Miami, FL

QUANTITY		Not available at publication

DISTRIBUTION		Nationwide

REASON		The amount of effort necessary to remove the adhesive liner from the patch.

-----------------------------------------------------------------------------------------------------------------------------------------------