Procedure and revision number: SN11269CLSI Rev 1/06

Procedure:  Premier Platinum HpSA PLUS

Meridian Catalog #601396

Institution:

Address:

Department:

| Prepared By: | Date Adopted | Supercedes procedure #: |
| --- | --- | --- |

| Distributed To: | # of Copies | Distributed To: | # of Copies |
| --- | --- | --- | --- |

NOTE: This procedure is provided to Meridian’s customers to assist with the development of laboratory procedures.  This document was derived from, and was current with, the instructions for use (IFU) that accompanies the product at the time it was created. The user is instructed to consult the IFU packaged with the product to ensure currency of the procedure prior to adapting the document to routine laboratory use and periodically thereafter to ensure future IFU modifications, which might effect this procedure, are identified.   Any modifications to this document are the sole responsibility of the person making the modifications.

PRINCIPLE:

The Premier Platinum HpSA PLUS enzyme immunoassay (EIA) is an in vitro qualitative procedure for the detection of Helicobacter pylori antigens in human stool.  Test results are intended to aid in the diagnosis of H. pylori infection and to monitor response during and post-therapy in patients.  Accepted medical practice recommends that testing by any current method, to confirm eradication, be done at least four weeks following completion of therapy.

Premier Platinum HpSA PLUS is a microwell-based enzyme immunoassay that detects H. pylori antigens present in human stool. No calculations are required and the visual color change makes the interpretation of results objective and simple. In addition, the HpSA test permits assessment of established or novel anti-H. pylori treatment during and post-therapy to monitor for treatment effectiveness, relapse or eradication.  Premier Platinum HpSA PLUS is a modification of Premier Platinum HpSA that provides increased signal strengths with positive test results and better discrimination between low positive and negative tests.

The Premier Platinum HpSA PLUS test utilizes a plurality of monoclonal anti-H. pylori capture antibodies adsorbed to microwells. (Plurality is defined as a mixture of monoclonal antibodies.)  Diluted patient samples and a conjugate(peroxidase conjugated to a plurality of monoclonal antibodies) are added to the wells and incubated for one hour at room temperature.  A wash is performed to remove unbound material.  Substrate is added and incubated for ten minutes at room temperature.  Color develops in the presence of bound enzyme.  Stop Solution is added and the results are interpreted visually or spectrophotometrically.

SPECIMEN:

Preferred sample type and volumes:

Liquid or semi-solid stool:  200-300 uL (100 uL for testing and he remaining for retesting as needed)

Formed solid stool:  20 mm sample (5-8 mm for testing and the remaining for retesting if needed)

Collection containers:  airtight transport containers

See section on preparation

Undesirable specimen characteristics:

- Watery stool specimens

- Stool in transport media, swabs, or preservatives are inappropriate for testing

Interfering substances:

The following substances, that may be present in human stool do not interfere with positive or negative test results at the stated concentrations per 500 uL human stool:  TUMS - 10 mg, Mylanta - 0.84 mg, Pepto Bismol - 0.35 mg, Tagamet - 1 mg, Prilosec OTC - 1 mg, barium sulfate - 10 mg, whole blood - 100 uL,  mucin - 6.7 mg, human hemoglobin (ie, dark stool) – 15 mg, steric + palmitic acids (ie, fatty stool) – 7.9 mg.

Preparation:

- No special patient preparation is required.

- Specimen collection:

  - Collect stool specimens into a clean, airtight container with no preservative.

  - All stool specimens should be transported in an airtight transport container and stored at 2-8 C until tested. The specimen should be tested as soon as possible, but may be held up to 72 hours at 2-8 C prior to testing.  (See SPECIMEN PREPARATION section for instructions on diluting samples.)

  - If testing cannot be performed within this time frame, specimens should be frozen immediately upon receipt and stored frozen (-20 C to –80 C) until tested.  Specimens may be frozen and thawed twice.

  - Use only the Sample Diluent provided with the kit for diluting specimens.

- Specimen preparation:

  - Using a pipetting device, add 500 µL of Sample Diluent to a clean test tube.

  - Mix stool as thoroughly as possible prior to pipeting.

  - For liquid or semi-solid stools - Using the supplied transfer pipette, add 100 µL (second mark from the tip of the pipette) of stool into Sample Diluent.  Using same pipette, gently withdraw and expel the stool suspension several times, then vortex 15 seconds.  Save the transfer pipette in the sample for later use.

  - For formed/solid stools - Using a wooden applicator stick, transfer a small portion (5-6 mm diameter) of thoroughly mixed stool into Sample Diluent.  Emulsify stool using the wooden applicator stick, then vortex 15 seconds.

  - Stool specimens may be centrifuged after dilution.  Centrifuge at approximately 2750 x G for 5 minutes or until solid matter separates from liquid.  Proceed with the assay after recovering supernate.

MATERIALS AND EQUIPMENT:

Materials:

    - Antibody Coated Microwells - Breakaway plastic wells coated with a plurality of murine monoclonal antibodies specific for H. pylori.

    - Positive Control - Inactivated H. pylori at approximately 37.5 µg protein/mL in a pH 7.2, 10 mM phosphate-buffered solution with 0.02% Thimerosal.

    - Sample Diluent/Negative Control - pH 7.2, 10 mM phosphate-buffered solution with 0.02% Thimerosal.

    - Premier 20X Wash Buffer - pH 6.8, 180 mM phosphate-buffered solution with 0.2% Thimerosal.

    - Enzyme Conjugate – A plurality of murine monoclonal antibodies specific for H. pylori conjugated to horseradish peroxidase in a pH 7.8, 50 mM Tris-buffered solution containing 0.02% Thimerosal.

    - Premier Substrate I - a-buffered solution containing urea peroxide and tetramethylbenzidine. (pH 5.0)

    - Premier Stop Solution I – 1M phosphoric acid. CAUTION: Avoid contact with skin. Flush with water if contact occurs.

    - Distilled or deionized water

Equipment:

a.	Transfer pipettes (one for each test sample).  Each pipette is marked to indicate 50 µL, 100 µL, 200 µL and 300 µL volumes.

- Microwell strip holder

- Microwell strip sealer

- Wooden stick applicators

- Test tubes (12 x 75 mm) for dilution of sample

- Squirt bottle

- Graduated cylinder for making 1X Wash Buffer

- Validated EIA plate reader capable of reading absorbance at 450 or 450/630 nm*

- Validated, semiautomated plate washer (e.g., BioTek Elx50) *

Reagent preparation:

    - Bring the entire kit, including microwell pouch, to 19-27 C before use.

    - Prepare 1X Wash Buffer as needed. For example: 4.0 mL of Premier 20X Wash Buffer + 76.0 mL of distilled or deionized water is sufficient to wash one strip.  Place in a clean squirt bottle.  The 1X Wash Buffer can be stored at 19-27 C for up to three months.

Performance Considerations:

- All reagents are for in vitro use only.

- Patient specimens may contain infectious agents and should be handled and disposed of as potential biohazards.

- All reagents should be mixed gently before use.

- Do not interchange the Microwells, Enzyme Conjugate, Substrate Reagent, or Positive Control reagents between lots.  (The Sample Diluent, Premier 20X Wash Buffer and Premier Stop Solution I are interchangeable provided the reagents are within their assigned expiration dates when used.)

- Allow reagents to warm to 19-27 C before use.

- Hold reagent vials vertically at suitable distance above the well to insure proper drop size and delivery.

- Do not use kit components beyond labeled expiration date.

- Replace colored caps on correct vials.

- Dispose of used wash buffer and all test materials in an appropriate container.  Treat waste as a potential biohazard.

- The Positive Control reagent contains inactivated H. pylori.  It should be handled, however, as a potential biohazard.

- Avoid skin contact with Premier Stop Solution I (1M phosphoric acid).  Flush with water immediately if contact occurs.

- Do not reuse microwells.

- Unused microwells must be placed back inside resealable pouch.  It is important to protect strips from moisture.

- The transfer pipettes provided with this kit must be used for specimen preparation and transfer.  Use one per specimen.

- Avoid splashing when dispensing diluted stool into microwells by placing the transfer pipette tip about halfway down the well and dispensing slowly down the side of well.

- Microwell washing is to be performed precisely as directed in assay procedure.  Inadequate washing may be the cause of elevated background in any EIA protocol.

- All reagents except the Premier 20X Wash Buffer are provided already diluted to the proper concentration.

- Any deviation below or above set incubation times may affect sensitivity and specificity and should be avoided.

- Stool must be mixed thoroughly (regardless of consistency) to insure a representative sample prior to pipetting.

- Some precipitation may occur in 20X Wash Buffer when it is stored at 2-8 C.  The precipitate will dissolve when a working dilution is made with the Wash Buffer.

- Do not use vials that lack a label, a lot number, or an expiration date.

Storage Requirements:

    - Store all kit reagents as 2-8 C.  Return the kit to the refrigerator promptly after use.

    - Keep microwells in their pouch until the pouch reaches room temperature to avoid condensation.  Return unused microwells and the desiccant to the ziplock pouch and seal tightly.

    - Diluted 1X Wash Buffer may be stored at 19-27 C and used for up to three months.

CALIBRATION:

Not applicable to this assay.

QUALITY CONTROL:

Daily:

      - At the time of each use, kit components should be visually examined for obvious signs of microbial contamination, freezing or leakage.  Do not use contaminated or suspect reagents.

      - The Positive and Negative Controls must be used with each test batch.  See the INTERPRETATION OF RESULTS section above for a description of the expected results for control reagents.

      - Tests should be considered invalid when either control reagent does not produce the expected results.  In such cases, repeat tests and controls.  If, on repeat testing, the expected reactions are still not observed and the reagents are still within their expiration date, contact Meridian’s Technical Support Center at 1-800-343-3858 for assistance.

Procedural controls:

1.	The controls are used to monitor reagent reactivity.  Failure of the controls to produce the expected results can mean that one or more of the reagents are no longer reactive at the time of use, the test was not performed correctly, or that reagents or samples were not added.

2.	The Positive Control will not ensure precision at the cut-off.

External Positive and Negative Controls:

1. 	The Positive Control and Sample Diluent/Negative Control reagents provided as components of the assay serve the purpose of external control reagents.

  - Specimen matrix interference has not been observed in this assay as samples are significantly diluted before testing in Sample Diluent.  For this reason, the positive and negative control reagents supplied as part of this assay are prepared in the matrix of the Sample Diluent.  If control materials that are identical in composition to test samples are preferred, the user can prepare these by diluting known positive and negative specimens in Sample Diluent according to the SPECIMEN PREPARATION section of this insert.   Add 100 µL of user prepared controls to test wells.

CAUTION:  Do not report patient test results when control results deviate from the expected.

TEST PROCEDURE:

  - After the pouch has reached temperature, break off the required number of microwells (one well for each specimen, plus one positive and one negative control well per batch). Place the microwells in the microwell strip holder and record the location of all wells.  Unused microwells must be resealed in the pouch immediately.

  - Using the specimen transfer pipette, add 100 µL of diluted stool (second mark from the tip of the pipette) to the appropriate well. (Place the pipette tip halfway into well and let the sample slowly run down side of well.)

  - Add two free falling drops of Positive Control and 100 µL of Sample Diluent /Negative Control to the appropriate wells.

  - Add one free falling drop (approximately 50 µL) of Enzyme Conjugate to each well.  Firmly shake/swirl the plate for 30 seconds.

  - Cut plate sealer to size and press firmly onto top of microwells to seal.  Incubate the plate for one hour at 19-27 C.

  - Carefully remove the plate sealer and wash wells:

    - Manual method:

      - Dump plate contents firmly into a biohazard receptacle.

      - Bang the inverted plate on a clean stack of paper towels.

      - Fill all wells with 1X Wash Buffer, directing stream of buffer to the sides of the wells to avoid foaming.

      - Repeat wash cycle (dump, bang on fresh towels, fill) four times for a total of five wash cycles.  After the last fill, dump and bang plates on fresh towels hard enough to remove as much excess wash buffer as possible, but do not allow wells to completely dry at any time.

    - Semiautomated method using validated equipment

      - Aspirate the contents of the well.

      - Fill the wells to the top (approx 300-350 µL/well) with 1X Wash Buffer then aspirate.  The washer manifold should be adjusted to ensure no foaming occurs during the filling of the wells and that the wells are thoroughly aspirated after each wash.

      - Repeat step ii a minimum of four more times.  Following the last wash, test wells should be thoroughly aspirated to remove as much moisture as possible.

  - Clean the underside of all wells with a lint-free tissue.

  - Add two free falling drops (approx. 100 µL) of Premier Substrate Solution I to each well.  Firmly shake/swirl the plate for 30 seconds.  Incubate for ten minutes at 19-27 C.

  - Add two free-falling drops (approx. 100 µL) of Premier Stop Solution I to each well.  Firmly shake/swirl the plate for 30 seconds.

Note: Initial color of positive reaction is blue, which changes to yellow upon addition of Premier Stop Solution I.

  - Inspect and record reactions.  Test results can be read visually or using a spectrophotometric reader.

    - Visual Determination - Read within 15 minutes after adding Premier Stop Solution I.

    - Spectrophotometric Determination - Zero EIA reader on air.  Wipe underside of wells with a lint-free tissue.  Read absorbance at 450 nm or 450/630 nm within 15 minutes of adding Premier Stop Solution I.

CALCULATIONS:

There are no calculations associated with the procedure.

INTERPRETATION OF RESULTS;

The following interpretations apply to both initial diagnosis and monitoring of anti- H. pylori therapy.

Visual Reading

Negative = colorless to faint yellow

Positive   = definite yellow color

To be called positive, a faint yellow color must be confirmed by a spectrophotometric reading.  If a spectrophotometer is not available, the cut-off must be determined by an alternative method.

Spectrophotometric Single Wavelength (450 nm)

Negative:  < 0.140

Positive:  ( 0.140

Negative Control: < 0.140

Positive Control: ( 0.640

Spectrophotometric Dual Wavelength (450/630 nm)

Negative: < 0.100

Positive:  ( 0.100

Negative Control: < 0.100

Positive Control: ( 0.600

If a Negative Control is < 0.000, reblank the plate reader to air and reread the plate.

REPORTING OF RESULTS:

A positive result indicates the presence of H. pylori antigens.  A negative result indicates the absence of H. pylori antigens, or that the level of antigens is below what can be detected by the assay.    The magnitude of the OD above the cut-off is not indicative of the severity or extent of H. pylori infection, nor can it be correlated to an endpoint titer.  Extremely strong positive reactions may yield a purple precipitate within a few minutes of stopping the reaction.

| Result Obtained | Suggested Report |
| --- | --- |
| Positive | Positive for H. pylori infection |
| Negative | Negative for H. pylori infection |
| Invalid | DO NOT REPORT ANY RESULT. REPEAT TEST. |

PROCEDURE NOTES:

1.	Studies on the epidemiology of H. pylori have shown that this organism is present worldwide. 18,23,24 Gastritis caused by H. pylori has been shown to correlate with age, ethnic background, family size and socioeconomic class. 25, 26   The prevalence of H. pylori infection in a given population can vary from 20% to 90%. In patients diagnosed with duodenal ulcers, however, it has been shown in every age group to be approximately 80%. 18 Currently recommended eradication treatments have an efficacy rate between 75% and 90%.

2.	The Premier Platinum HpSA PLUS test detects the presence of H. pylori antigens in human stool.  Expected values for a given population should be determined for each laboratory. The rate of positivity may vary depending on geographic location, method of specimen collection, handling and transportation, test employed and general health environment of patient population under study.  As demonstrated by Premier Platinum HpSA in tests conducted in the United States, Canada and Italy, incidence of disease ranged from 34% to 53% to 69% respectively.

3.	Since its discovery over 20 years ago by Marshall and Warren, 1 Helicobacter pylori is now recognized as one of the most common and medically important pathogens worldwide. 2 Helicobacter pylori has been firmly established as an etiologic agent in chronic gastritis and peptic ulcer disease, and has been associated with mucosa-associated lymphoid tissue lymphoma and gastric adenocarcinoma. 1,3-7

The ecological niche in humans appears to be restricted to the stomach and the duodenum.  Patients who harbor the organism are divided into two basic groups.  The first group shows no signs or symptoms of gastrointestinal disease and is considered as “colonized”.  The second group shows gastrointestinal signs and symptoms and is considered as “infected”.  The process by which an individual becomes colonized or infected is still under investigation. 3,4,8-10  Many possible routes of transmission of Helicobacter pylori to humans such as animals, contaminated water and oral reservoirs have been suggested. 11

Diagnostic tests for H. pylori can be categorized as invasive (endoscopy, biopsy) or noninvasive (serology, urea breath test and stool antigen test).  In invasive testing, a biopsy is taken from the upper gastrointestinal tract and examined microscopically.  The tissue is also cultured for H. pylori or evaluated in the rapid urease test.  This strategy offers the advantages of detecting an active infection and has high specificity and a high positive predictive value.  The disadvantages of invasive testing include risk and discomfort to the patient and colonization in patches that might be missed by biopsy.  Culture of biopsy material is time consuming and can yield false-negative results due to inherent technical difficulties. 3,12-15

The urea breath test (UBT) is a type of noninvasive test that detects the highly active urease of H. pylori.  Although UBT is highly sensitive and specific, it has a number of significant drawbacks.  UBT is time consuming, requires specialized detection equipment and involves the ingestion of isotopically labeled urea by the patient. 3,16,17  Serological tests, also noninvasive, based on the detection of IgG against H. pylori are useful for primary screening of patients that present with uncomplicated infections, yet they do not distinguish between past exposure and active infection. 6,11,18  The stool antigen test has been evaluated extensively and has been accepted as an accurate non-invasive test both before and after treatment. 19-21  The recent Maastricht 2 Consensus Report recommends the use of the stool antigen and UBT tests as an aid in the diagnosis of H. pylori disease in the primary care setting. 22

PERFORMANCE CHARACTERISTICS:

Clinical evaluations performed with the first generation Premier Platinum HpSA demonstrated that an ELISA-based assay could reliably and predictably detect H. pylori antigen in human stool in symptomatic patients.  Studies also demonstrated the test can be used to monitor the efficacy of eradication therapy.

The Premier Platinum HpSA was evaluated on 200 symptomatic adults at one Midwestern United States location, one site in Canada, and two sites in Italy.  The patients studied had a wide cross-section of gastric pathologies noted, including: antral gastritis (n=81), antral gastropathy (n=25), antral erosions (n=24), esophagitis (n=21), duodenal ulcer (n=15), erosive duodenitis (n=10), GERD (n=10), “normal” (n=10), duodenitis (n=9), gastric ulcer (n=8), total stomach gastritis (n=6), hiatal hernia (n=6), Schatzki’s ring (n=4), pyloric ulcer (n=2), and esophageal ulcer (n=1).  HpSA test results were compared to diagnosis of H. pylori infection as judged by objective reference methods (culture, rapid urease, histology and UBT).  Patients were considered positive if culture was positive, or if two or more of the other three tests were positive.  Nine patients with negative or no culture results, and only one other test positive, were considered unevaluable.  The HpSA test was 96.1% sensitive, 95.7% specific and showed 95.9% correlation with H.pylori infection.   Confidence intervals were calculated by the exact binomial method.

Trial Site #1

| TEST | DIAGNOSIS | Sensitivity | Specificity | Pos. PV | Neg. PV | Correlation |
| --- | --- | --- | --- | --- | --- | --- |
| Method | Result | Infected | Not Infected | ±95% CI | ±95% CI | ±95% CI | ±95% CI | ±95% CI |
| PPHpSA | Pos | 17 | 3 | 94.4% | 91.4% | 85.0% | 97.0% | 92.5% |
| EIA | Neg | 1 | 32 | 72.7-99.9% | 76.9-98.2% | 62.1-96.8% | 84.2-99.9% | 81.8-97.9% |
|  | Equ | 0 | 0 |  |  |  |  |  |

Reference Methods: Histology, Rapid Urease, Breath Test. Readings Single and Dual Wavelength.

Trial Site #2

| TEST | DIAGNOSIS | Sensitivity | Specificity | Pos. PV | Neg. PV | Correlation |
| --- | --- | --- | --- | --- | --- | --- |
| Method | Result | Infected | Not Infected | ±95% CI | ±95% CI | ±95% CI | ±95% CI | ±95% CI |
| PPHpSA | Pos | 9 | 0 | 100.0% | 100.0% | 100.0% | 100.0% | 100.0% |
| EIA | Neg | 0 | 8 | 66.4-100.0% | 63.1-100.0% | 66.4-100.0% | 63.1-100.0% | 80.5-100.0% |
|  | Equ | 0 | 0 |  |  |  |  |  |

Reference Methods: Histology, Rapid Urease, Culture, Breath Test. Readings Single and Dual Wavelength.

Trial Site #3

| TEST | DIAGNOSIS | Sensitivity | Specificity | Pos. PV | Neg. PV | Correlation |
| --- | --- | --- | --- | --- | --- | --- |
| Method | Result | Infected | Not Infected | ±95% CI | ±95% CI | ±95% CI | ±95% CI | ±95% CI |
| PPHpSA | Pos | 44 | 0 | 97.8% | 100.0% | 100.0% | 96.0% | 98.6% |
| EIA | Neg | 1 | 24 | 88.2-99.9% | 85.8-100.0% | 92.0-100.0% | 79.6-99.9% | 92.2-100.0% |
|  | Equ | 1 | 0 |  |  |  |  |  |

Reference Methods: Histology, Rapid Urease, Culture, Breath Test. Readings Single Wavelength.

Trial Site #4

| TEST | DIAGNOSIS | Sensitivity | Specificity | Pos. PV | Neg. PV | Correlation |
| --- | --- | --- | --- | --- | --- | --- |
| Method | Result | Infected | Not Infected | ±95% CI | ±95% CI | ±95% CI | ±95% CI | ±95% CI |
| PPHpSA | Pos | 29 | 1 | 93.5% | 96.3% | 96.7% | 92.9% | 94.8% |
| EIA | Neg | 2 | 26 | 78.6-99.2% | 81.0-99.9% | 82.8-99.9% | 76.5-99.1% | 85.6-98.9% |
|  | Equ | 2 | 0 |  |  |  |  |  |

Reference Methods: Histology, Rapid Urease. Readings Dual Wavelength.

Combined Data From All Sites

| TEST | DIAGNOSIS | Sensitivity | Specificity | Pos. PV | Neg. PV | Correlation |
| --- | --- | --- | --- | --- | --- | --- |
| Method | Result | Infected | Not Infected | ±95% CI | ±95% CI | ±95% CI | ±95% CI | ±95% CI |
| PPHpSA | Pos | 99 | 4 | 96.1% | 95.7% | 96.1% | 95.7% | 95.9% |
| EIA | Neg | 4 | 90 | 90.4-98.9% | 89.5-98.8% | 90.4-98.9% | 89.5-98.8% | 92.2-98.2% |
|  | Equ | 3 | 0 |  |  |  |  |  |

Therapeutic Monitoring:

Four sites examined the utility of the stool antigen test for monitoring anti- H. pylori treatment in 97 patients who initially tested positive by endoscopy (culture, histology, and rapid urease).  Premier Platinum HpSA testing and endoscopic biopsies were performed 4 weeks after completion of physician-prescribed, H. pylori eradication therapy. The test results are compared in the following table.  Culture, histology, and rapid urease were used to determine eradication as defined by FDA guidelines. 22

| Overall: HpSA vs. 4 Week Scope |
| --- |
| HpSA | 4 Weeks Post Treatment |
| Result | Infected | Eradicated |
| Positive | 18 | 3 |
| Negative | 1 | 73 |
| Statistic | Value | 95% CI |
| Sensitivity | 94.7% | 74.0-99.9% |
| Specificity | 96.1% | 88.9-99.2% |
| Positive PV | 85.7% | 63.7-97.0% |
| Negative PV | 98.6% | 92.7-100% |
| Correlation | 95.8% | 89.6-98.8% |

Premier Platinum HpSA correctly identified 18/19 (94.7%) of the infected and 73/76 (96.1%) of the eradicated patients.  Two of the 97 stools were equivocal by HpSA (2%). The false negative stool was from a patient that was positive by culture, histology and rapid urease.  Three false positive HpSA results were obtained from patients that were negative by all other methods (culture, histology, and rapid urease).

Response to treatment is generally noted by a negative HpSA test within 5 to 7 days after initiating treatment.  Positive results at this time, or later, indicate ineffective therapy or recurrence.  Recurrence can result from lack of patient compliance with the drug regimen, ineffective drugs, resistant strains of H. pylori, improper dosage, etc.  Recurrent H. pylori infection generally occurs by four weeks after termination of therapy.  Occasionally, however, infections will remain cryptic beyond four weeks.  This observation supports accepted medical practice that determination of eradication utilizing any diagnostic method should be done at least four weeks following completion of therapy.  The figures below are typical response profiles for successful and unsuccessful eradication therapies.  The vertical bar indicates completion of therapy (Day 0). Days to the left of Day 0 reflect the period that patients were taking drugs.  The positive cut-off is shown as horizontal dashed lines.

### Eradicated Patient Profiles

### Non-Eradicated (Infected) Patient Profiles

Comparison of Premier Platinum HpSA PLUS to Premier Platinum HpSA:

Tests with 291 samples from symptomatic patients collected either prior to or following treatment were used to demonstrate that Premier Platinum HpSA PLUS performed similarly to Premier Platinum HpSA.   Thirty three of these samples were originally evaluated in an earlier trial to demonstrate the effectiveness of Premier Platinum HpSA.  Test performance including 95% confidence intervals is detailed in the following table.

| PP HpSA PLUS | PP HpSA (Predicate) |
| --- | --- |
|  | Positive | Negative | Indeterminate |
| Positive | 94 | 10 | 3 |
| Negative | 0 | 183 | 1 |
| Agreement | Positive Test | Negative Test | Overall |
|  | 94/94 = 100% | 183/193 = 94.8% | 277/287 = 96.5% |

Eight of the ten samples that were positive by Premier Platinum HpSA PLUS, but negative by Premier Platinum HpSA, were positive by CLO, histology or UBT testing.  The three samples that were positive by Premier Platinum HpSA PLUS but indeterminate by Premier Platinum HpSA were positive by CLO, histology or UBT testing.  The one sample, that was negative by Premier Platinum HpSA PLUS but indeterminate by Premier Platinum HpSA, was negative by CLO, histology or UBT testing.

Reproducibility

Assay precision, intra-assay variability and inter-assay variability were assessed with a reference panel prepared from high positive samples (n = 2), low negative samples (n = 2), and low positive and high negative specimens (n = 1 each).  The latter were diluted to near the assay limit of sensitivity.  Nine replicates each of the low positive and high negative samples were included in the panel to bring the total cohort to 22 reference specimens.  Each reference specimen was coded to prevent its identification during testing.  Each was evaluated twice per day for three consecutive days by three different laboratories.  High negative samples (OD value just below 0.100) produced weakly positive results (OD values just above 0.100) in 42 out of 162 tests.  It is expected that high negative samples prepared at the cut-off will produce weakly positive results 50% of the time.  (See EP12-A, User protocol for evaluation of qualitative performance; approved guideline; NCCLS/CLSI, Vol. 22, no.14, 2002.)  Low positive, high positive and low negative samples produced the correct results 100% of the time.  Reproducibility was 100% with no intra-assay and inter-assay variability for samples prepared above or below the limit of analytical sensitivity.

Assay specificity

The specificity of Premier Platinum HpSA PLUS was tested by utilizing the following bacterial, yeast or viral strains.  Positive and negative stools were spiked with ( 1.2 X 109 bacterial or yeast organisms/ml  and tested by Premier Platinum HpSA PLUS.  The concentration of viral organisms was not calculated.  None of the organisms affected positive or negative test results.

Microorganism or virus

| Adenovirus | Aeromonas hydrophila | Campylobacter lari | Campylobacter fetus |
| --- | --- | --- | --- |
| Campylobacter jejuni | Campylobacter jejuni 2 | Campylobacter jejuni solution | Candida albicans |
| Citrobacter freundii | Clostridium difficile | Clostridium perfringens | Enterobacter cloacae |
| Enterococcus faecalis | Escherichia coli 0157:H7 | Escherichia coli 8739 | Escherichia coli 9637 |
| Escherichia fergusonii | Escherichia hermannii | Escherichia hermannii EMDi-64 | Helicobacter pylori |
| Klebsiella pneumoniae | Lactobacillus lactis | Listeria monocytogenes | Peptostreptococcus anaerobius |
| Proteus vulgaris | Pseudomonas aeruginosa | Pseudomonas fluorescens | Salmonella Group B |
| Salmonella typhimurium | Serratia liquefaciens | Serratia marcescens | Shigella boydii |
| Shigella flexneri | Shigella dysenteriae | Shigella sonnei | Staphylococcus aureus |
| Staphylococcus aureus (Cowans 1) | Staphylococcus epidermidis | Streptococcus faecalis | Salmonella enterica serovar Hilversum |
| Salmonella enterica subsp. Enterica serovar Hilversum | Salmonella enterica subsp. Enterica serovar Minnesota | Yersinia enterocolitica | Rotavirus |

LIMITATIONS OF THE PROCEDURE:

1. 	The test is qualitative and no quantitative interpretation should be made with respect to the values.

2. 	Test results should be used in conjunction with information available from the patient clinical evaluation and other diagnostic procedures.

3. 	Antimicrobials, proton pump inhibitors and bismuth preparations are known to suppress H. pylori and ingestion of these prior to H. pylori testing (culture, histology, rapid urease, UBT, antigen) may give a false negative result.  If a negative result is obtained for a patient ingesting these compounds within two weeks prior to performing the Premier Platinum HpSA PLUS test, it may be a false-negative result and the test should be repeated on a new specimen obtained two weeks after discontinuing treatment.  A positive result for a patient ingesting these compounds, within two weeks prior to performing the Premier Platinum HpSA PLUS test, should be considered accurate.  As an example, patients with H. pylori were placed on a proton pump inhibitor (Lansoprazole) or bismuth for 2 weeks, and tested with the Premier Platinum HpSA and a urea breath test.  Patients were then taken off treatment for 2 weeks and retested.  At the end of treatment, both assays were negative for some patients, but returned to positive 2 weeks post-treatment (see table).

|  |  | Premier Platinum HpSA | Breath Test |
| --- | --- | --- | --- |
| Treatment | Time Point | Pos/Total | % Positive | Pos/Total | % Positive |
| PPI | End of Treatment | 15/20 | 75.0% | 12/20 | 60.0% |
|  | 2 Weeks Post-Treatment | 19/20 | 95.0% | 18/20 | 90.0% |
| Bismuth | End of Treatment | 15/20 | 75.0% | 11/20 | 55.0% |
|  | 2 Weeks Post-Treatment | 19/20 | 95.0% | 18/20 | 90.0% |

4. 	Performance characteristics have not been established for watery, diarrheal stools.

5. 	Performance characteristics have not been established in asymptomatic populations.

6. 	The Premier Platinum HpSA PLUS test can detect ( 4.67 ng H. pylori protein/mL of stool.  (Analytical sensitivity limit.)

7. 	H2 blockers do not interfere with positive results.

REFERENCES:

- Marshall BJ, Warren JR. Unidentified curved bacilli in the stomach of patients with gastritis and peptic ulceration. Lancet 1984;i: 1311-1314.

- Versalovic J. Helicobacter pylori. Pathology and diagnostic strategies. Am J Clin Pathol 2003;119:403-12.

- Dunn BE, Cohen H, Blaser MJ. Helicobacter pylori. Clin. Microbiol. Rev. 1997;10:720-41.

- Gregson DB, Simor AE. Helicobacter pylori and inflammatory disease of the stomach and duodenum. Stomach and Bowel, 1991;Sep:29-30, 35-7.

- Parsonnet J, Friedman GD, Vandersteen DP et al. 1991. Helicobacter pylori infection and the risk of gastric carcinoma. N Engl J Med 1991;325:1127-31.

- Sanders MK, Peura DA. Helicobacter pylori-associated diseases. Curr Gastro Reports. 2002;4:448-54.

- Uemura N, Okamoto S, Yamamoto S et al. Helicobacter pylori infection and the development of gastric cancer. N Engl J Med. 2001;345:784-9.

- Graham DY. I. Helicobacter pylori: Its epidemiology and its role in duodenal ulcer disease. J Gastroenterol Hepatol 1991;6:105-13.

- Malaty HM, Graham DY, Klein PD et al. Transmission of Helicobacter pylori infection. Studies in families of healthy individuals. Scand J Gastroenterol 1991;26:927-32.

- Marshall BJ, Armstrong JA, McGechie DB, Glancy RJ. Attempt to fulfill Koch’s postulates for pyloric campylobacter. Med J Austral 1985;142:436-439.

- Veralovic J, Fox JG. Helicobacter. In: Murray PR et al eds. Manual of clinical microbiology, 8th ed. Washington DC:ASM Press, 2003:915-28.

- Alpert LC, Graham DY, Evans DJ Jr et al. Diagnostic possibilities for Campylobacter pylori infection. Eur J Gastroenterol Hepatol 1989;1:17-26.

- Barthel JS, Everett ED. Diagnosis of Campylobacter pylori infections: The “gold standard” and the alternatives. Rev Infect Dis 1990;12:S107-14.

- Nichols L, Sughayer M, DeGirolami PC et al. Evaluation of diagnostic methods for Helicobacter pylori gastritis. Am J Clin Pathol 1991;95:769-73.

- Vaira D, Vakil N. Blood, urine, stool, breath, money and Helicobacter pylori. Gut 2001;48:287-9.

- Graham DY, Klein PD, Evans DJ Jr et al. Campylobacter pylori detected noninvasively by the 13C-urea breath test. Lancet 1987;i:1174-7.

- Graham DY, Malaty HM, Evans DG et al. Epidemiology of Helicobacter pylori in an asymptomatic population in the United States. Effect of age, race, and socioeconomic status. Gastroenterology 1991;100:1495-1501.

- Talley NJ, Newell DG, Ormand JE et al. Serodiagnosis of Helicobacter pylori Comparison of enzyme-linked immunosorbent assays. J Clin Microbiol 1991;29:1635-9.

- Vaira D, Malfertheiner P, Mégraud F et al. Diagnosis of Helicobacter pylori infection with a new non-invasive antigen-based assay. Lancet 1999;354:30-3.

- Vaira D, Malfertheiner P, Mégraud F et al. Non invasive antigen-based assay for assessing Helicobacter pylori eradication: a European Multicentre Study. Am J Gastrenterol 2002;95:925-9.

- Montiero L, de Mascarel A, Sarrasqueta A. et al. Diagnosis of Helicobacter pylori Infection: Noninvasive Methods Compared to Invasive Methods and Evaluation of Two New Tests. Am J Gastroenterol 2001;96:353-8.

- Malfertheiner P, Megraud F, Morain O et al. Current concepts in the management of Helicobacter pylori infection – The Maastricht 2-2000 Consensus Report. Aliment Pharmacol Ther 2002;16:167-80.

- Murray DM. Clinical relevance of infection by Helicobacter pylori. Clin Microbiol Newlett 1993;15:33-7.

- Loffeld RJLF, Stobberingh E, Van Spreeuwel JP et al. The prevalence of anti-Helicobacter (Campylobacter) pylori antibodies in patients and healthy blood donors. J Med Microbiol 1991;32:105-9.

- Morris A, Nicholson G, Lloyd G et al. Seroepidemiology of Campylobacter pyloridis. N. Z. Med. J. 1986;99:657-9.

- Fiedorek SC, Malaty HM, Evans D Letal. Factors influencing the epidemiology of Helicobacter pylori infection in children.  Pediatrics 1991;88:578-582.