## Intended Use

For the quantitative determination of albumin in serum on Hitachi analyzers.  For in vitro diagnostic use only.

Method History

Determination of serum albumin is usually made using an ultra centrifugation, salt fractionation, electrophoretic or dye binding method.  Dye binding procedures are the simplest to perform, and lend themselves to high volume testing and automation.  They are also the procedures most widely used in combination with total protein determinations to yield an A/G ratio.1,2  In 1953, the use of methyl orange3 for direct determination was described.  This method suffered from non-specific binding characteristics.4,5  The use of a HABA6 dye was introduced in 1954.  This method was specific for albumin but displayed poor sensitivity, poor correlation with electrophoresis methods and significant interference from bilirubin, lipids, salicylates, penicillin and sulfonamides.7

A bromcresol green (BCG) dye-binding procedure was first proposed in 1964.8 This procedure exhibited greater sensitivity and much lower susceptibility to interfering substances.  The original method has been optimized to improve correlation with electrophoretic methods.9 The present procedure follows a modification of the original BCG dye-binding procedure.

Several publications of the late 1970’s 10,11,12,13 reported that abnormal proteins will bind with BCG after the first minute.  The present procedure includes a reduced measuring time to eliminate abnormal globulin interference and offers linearity to 6.0 g/dl.

Principle

Albumin is bound by the BCG dye to produce an increase in blue-green color measured at 600nm.  The color increase is proportional to the concentration of albumin present.

Reagents

Bromcresol green (BCG) 0.15 g/L, buffer, pH 4.20±0.1, surfactant, non-reactive ingredients and stabilizers.

Precautions

- This reagent is for in vitro diagnostic use only.

- Avoid ingestion.

- Avoid contact.  Reagent is an acid solution.  Flush with water when contact occurs.

- Reagent contains Sodium Azide as a preservative.  This may react with copper or lead plumbing to form explosive metal azides.  Upon disposal, flush with large amounts of water to prevent azide build up.

Reagent Preparation

Reagent is in a “ready to use” state.

Reagent Storage

Store reagent at room temperature.

Reagent Deterioration

The reagent should be a clear, yellow-green solution.  Turbidity or precipitation makes the reagent unsatisfactory and it should be discarded.

Specimen Collection and Storage 14

- Serum is the specimen of choice.

- Avoid excessive hemolysis since every 100 mg/dl of hemoglobin corresponds to about 100 mg/dl of albumin.

- Albumin in serum is reported stable for one week at room temperature (18-30°C) and approximately one month when stored in the refrigerator (2-8°C) and protected against evaporation.

Interferences

- See Young et al15 for a list of interfering substances.

- Ampicillin has been found to seriously interfere with BCG methods.16

Materials Provided

Albumin reagent

Materials Required but not Provided

- Controls

- Calibrator

- Hitachi analyzer

- Application and Instrument manuals

Procedure (Automated)(Hitachi 717)

TEST NAME		[ALB]

ASSAY CODE		[1-POINT]:[8]-[0]

SAMPLE VOLUME		[3] [2]

R1 VOLUME		[350] [100] [NO]

R2 VOLUME

WAVELENGTH		[700] [600]

CALIBRATION		[LINEAR] [0] [0]

STD (1)  CONC-POS		[0.0] [1]

STD (2)  CONC-POS 		[*] [*]

STD (3)  CONC-POS

STD (4)  CONC-POS

STD (5)  CONC-POS

STD (6)  CONC-POS

SD LIMIT			[0.1]

DUPLICATE LIMIT		[200]

SENSITIVITY LIMIT		[1000]

ABS LIMIT (INC/DEC)	[0] [INCREASE]

PROZONE LIMIT		[0] [LOWER]

EXPECTED VALUE		[*] [*]

PANIC VALUE		[*] [*]

INSTRUMENT FACTOR	[1.0]

For other Hitachi instrument applications, please contact the manufacturer’s technical service department.

Calibration

Follow instrument manual instructions for calibration. Refer to instrument manual instructions for calibration procedures and frequency.   It is recommended that each laboratory determine its own frequency of calibration.

Quality Control

The validity of the reaction should be monitored by use of control sera with known normal and abnormal albumin values.  These controls should be run at least with every working shift in which albumin assays are performed.  It is recommended that each laboratory establish its own frequency of control determination.

Calculation

Abs. = Absorbance

Abs. of Unknown   x  Conc. of  = Albumin (g/dl)

Abs. of Standard           Std.

Example:  If the Absorbance of the Unknown = 0.200 and the Absorbance of the Standard is 0.19 and the Standard Concentration = 3.5, then:

  - x  3.5  =  3.68 g/dl

0.190

Limitations

- The dye-binding properties of albumin, other than human, differ among species.17

- Samples with values above 6.0 g/dl should be diluted with 0.9% saline 1:1, re-run, and results multiplied by 2.  Samples with results below 0.5 g/dl should be done electrophoretically.

Expected Values1

3.5 – 5.3 g/dl

It is strongly recommended that each laboratory establish its own normal range.

Performance

- Linearity:  0.5 – 6.0 g/dl

- Comparison:   A comparison study performed between this method and another BCG method resulted in a correlation coefficient of 0.958 with a regression equation of y = 0.999 x-0.19.

- Precision:

Within Day		            Day to Day

Mean       S.D.	     C.V.%		Mean	S.D.	C.V.%

3.69	       0.02	     0.54		3.90	0.12	3.1

4.60         0.05	     1.07		4.80	0.10	2.1

References

- Tietz, N., Fundamentals of Clinical Chemistry, Philadelphia, W.B. Saunders, pp. 335-337 (1976).

- Davidson, I., Henry, J., Todd-Stanford Clinical Diagnosis by Laboratory Methods, Philadelphia, W.B. Saunders, p. 814 (1974).

- Bracken, J.S., Klotz, I.M., Am. J. Clin. Path. 23:1055 (1953).

- Lundh, B., Scand. J. Clin. Lab. Invest. 17:503 (1965).

- Rosenberg, R.M., et al. J. Am. Chem. Soc. 77:6502 (1955).

- Rutstein, D.D., et al, J. Clin. Invest 33:211 (1954).

- Arvan, D.A., Ritz, A., Clin. Chim. Acta. 26:505 (1969).

- Bartholomew, R., Delany, A., Proc. Australian Assoc. Clin. Biochem. 1:64 (1964).

- Dow, D., Pinto, PVC, Clin. Chem. 15:1006 (1969).

- Savory, J., et al, Clin. Chem. 22:1102 (1976).

- Corcoran, R., Duran, S., Clin. Chem. 23:765 (1977).

- Webster, D., Clin. Chem. 23:663 (1977).

- Gustaffson, J., Clin. Chem. 24:369 (1978).

- Doumas, B.T., Biggs, H.G., Standard Methods of Clinical Chemistry, Academic Press, N.Y., vol. 7, p. 175 (1972).

- Young D.S., et al, Clin. Chem. 21:1D (1975).

- Beng, C.G., Lim, K.L., Am. J., Clin. Path. 59:14 (1973).

- Spencer, D., et al, Anal. Clin. Biochem. 14:105 (1977).

Rev. 5/02  P803-HA901-01