QA-ES

User & Service Manual

QA-ES Electrosurgical Analyzer

P/N 14025

Copyright  ( 2001 by METRON. All rights reserved.

METRON:

USA				      FRANCE	________________	NORWAY________________

1345 Monroe NW, Suite 255A	      30, rue Paul Claudel	Travbaneveien 1

Grand Rapids, MI 49505	      91000 Evry, France	N-7044 Trondheim, Norway

Phone:	(+1) 888 863-8766	      Phone:   (+33) 1 6078 8899	Phone:	(+47) 7382 8500

Fax: 		(+1) 616 454-3350	      Fax:	    (+33) 1 6078 6839	Fax:	(+47) 7391 7009

E-mail:     support.us@metron-biomed.com    E-mail:   info@metron.fr 	    		        E-mail: support@metron.no

Disclaimer

METRON provides this publication as is without warranty of any kind, either express or implied, including but not limited to the implied warranties of merchantability or fitness for any particular purpose. Further, METRON reserves the right to revise this publication and to make changes from time to time to the content hereof, without obligation to METRON or its local representatives to notify any person of such revision or changes. Some jurisdictions do not allow disclaimers of expressed or implied warranties in certain transactions; therefore, this statement may not apply to you.

Limited Warranty

METRON warrants that the QA-ES Electrosurgical Analyzer will substantially conform to published specifications and to the documentation, provided that it is used for the purpose for which it was designed. METRON will, for a period of twelve (12) months from date of purchase, replace or repair any defective system, if the fault is due to a manufacturing defect. In no event will METRON or its local representatives be liable for direct, indirect, special, incidental, or consequential damages arising out of the use of or inability to use the QA-ES Electrosurgical Analyzer, even if advised of the possibility of such damages. METRON or its local representatives are not responsible for any costs, loss of profits, loss of data, or claims by third parties due to use of, or inability to use the QA-ES Electrosurgical Analyzer. Neither METRON nor its local representatives will accept, nor be bound by any other form of guarantee concerning the QA-ES Electrosurgical Analyzer other than this guarantee. Some jurisdictions do not allow disclaimers of expressed or implied warranties in certain transactions; therefore, this statement may not apply to you.

Table of Contents

1.	Introduction	1-1

1.1	QA-ES Description	1-1

1.2	QA-ES Specifications	1-1

1.3	General Information	1-2

2.	Installation	2-1

2.1	Receipt, Inspection  and Return	2-1

2.2	Setup	2-2

2.3	PRO-Soft QA-ES	2-2

3.	Operating QA-ES	3-1

3.1	Control Switches  and Connections	3-1

3.2	QA-ES Menu and  Function Keys	3-3

3.3	LCD Display Menu/  Messages (Overview)	3-3

3.4	LCD Display Menu/  Messages (Detail)	3-4

3.5	Printout	3-6

3.6	Foot Switch Output	3-6

3.7	Main Switch On/Off	3-6

4.	ESU Tests with QA-ES	4-1

4.1	Power Distribution	4-1

4.2	HF Current Leakage	4-1

4.3	REM Alarm	4-3

5.	Control and Calibration	5-1

5.1	Required Test Equipment	5-1

5.2        Preparation	5-1

5.3	Function Testing	5-2

5.4	Calibration	5-2

6.	Component Functions and Parts	6-1

6.1	Processor Board	6-1

6.2	Sampling Unit	6-3

6.3	Load Board	6-5

6.4	Component Parts	6-7

Appendix A – Diagrams 	A-1

Appendix B – Error Report Form	B-1

Appendix C – Improvement Suggestion Form	C-1

This page intentionally left blank.

Manual Revision Record

This record page is for recording revisions to your QA-ES User & Service Manual that have been published by METRON AS or its authorized representatives. We recommend that only the management or facility representative authorized to process changes and revisions to publications:

- make the pen changes or insert the revised pages;

- ensure that obsolete pages are withdrawn and either disposed of immediately, or marked as superseded and placed in a superseded document file, and;

- enter the information below reflecting that the revisions have been entered.

| Rev No | Date Entered | Reason | Signature of Person Entering Change |
| --- | --- | --- | --- |
| 1.30-1 | 4-30-01 | General Update |  |

This page intentionally left blank.

# Introduction

This chapter describes the METRON QA-ES Electrosurgical Analyzer, including its features and specifications.

## QA-ES Description

The METRON QA-ES Electrosurgical Analyzer (QA-ES) is a precision instrument designed to perform tests on high-frequency electrosurgical units (ESU) in accordance with national and international   standards, and is designed to be used by trained service technicians. Tests include:

- automatic power distribution measurement;

- crest factor measurement;

- RF leak measurement, and;

- return electrode monitor (REM) test

Testing is accomplished by measuring the ESU output against test loads that are set and adjusted in the QA-ES.  The QA-ES  can automatically execute a power distribution test with a load resistance ranging from 10 ohms to a maximum of 5200 ohms. The automatic measuring of the QA-ES, comprising crest factor measurements with a bandwidth of 10 MHz, ensures that the test result is reliable and reproducible.

Test results, shown in the QA-ES’s LCD display, can be printed out directly, or transferred to a PC via the PRO-Soft QA-ES test automation software. PRO-Soft lets you design test protocols, remotely control the QA-ES, and store the test results.

## QA-ES Specifications

| Generator Output: | RF LEAKAGE: From active electrode or neutral plate with an open or closed load circuit. |
| --- | --- |
| Mode Of Operation: | Manual or user-programmable. Can be remotely controlled with PC utilizing accessory PRO-Soft QA-ES software and RS-232 communication cable connection. |
| Measurements: | True RMS value of applied waveform. |
| RMS Bandwidth: | 30 Hz to 10 MHz (+3 dB). |
| Low Frequency Filter: | 100 Hz filter to avoid low frequency disturbance and/or interference. |
| Current: | 20 mA to 2200 mA. |

| Current Accuracy: | 20 - 2200 mA ± 2% of reading. |
| --- | --- |
| Load Resistance: | 10 - 2500 ohms in steps of 25 ohms (@ dc). 2600 - 5200 ohms in steps of 100 ohms (@ dc). |
| Additional Fixed Load: | 200 ohms, 500 watt maximum. |
| Crest Factor: | The higher of the two peak-measurements is used for calculation. |
| Range: | 1.4 - 16 (V peak voltage / V RMS). |
| Foot Switch Output: | The output triggers the measurement after a programmed delay time, defined as the time period from the activation of the foot switch to the beginning of data processing. The delay time is 200 ms - 4000 ms. |
| Peak To Peak Voltage: | 0 to 10 kV (closed load only) ± 10%. Measurement is taken between the active and dispersive electrodes with closed load only. |
| Oscilloscope Output: | 5 V/A uncalibrated, 100 mA RF current minimum input. |
| Isolation: | 10 kV isolation between measurement device and enclosure. |

## General Information

| Temperature Requirements: |
| --- |
| +15(C to +35(C when operating |
| 0(C to +50(C in storage |
| Display: |  |
| Type | LCD graphic display |
| Alphanumeric format | 8 lines, 40 characters |
| Graphics mode: | 240 x 64 point matrix |
| Display control: | 5 F-keys, enter, cancel and an encoder |
| Data Input/ Output (2): | Parallel printer port (1); Bi-directional RS -232C (1) for Computer control |
| Power Source: | From 115 VAC to 230 VAC, 48/66 MHz. |
| Mechanical Specifications: |
| Housing | Metal case |
| Height | 13.2 cm / 4.48 in. |  |
| Width | 34.2 cm / 11.61 in. |  |
| Length | 39.5 cm / 13.41 in. |  |
| Weight | 9.8 kg / 21.6 lbs. |  |
| Standard Accessories: |
| QA-ES Electrosurgical Analyzer (P.N. 14010) |
| Power Cord (P.N. 14300) |
| QA-ES User and Service Manual (P.N. 14025) |
| Additional Accessories: |
| E-Input Measuring Cable - Black (P.N. 11451) |
| E-Input Measuring Cable - Red (P.N. 11452) |
| Alligator Clamp - Black (P.N. 11461) |
| Alligator Clamp - Red (P.N. 11462) |
| Carrying case (P.N. 14100) |
| PRO-Soft QA-ES software (P.N. 12200) |
| PRO-Soft QA-ES DEMO (P.N. 14201) |
| PRO-Soft QA-ES User Manual (P.N. 14225) |
| Storage: |  |
| Store in the carrying case in dry surroundings within the temperature range specified. There are no other storage requirements. |

| Periodic Inspection: |  |
| --- | --- |
| The unit should be calibrated every 12 months. |

This page intentionally left blank

# Installation

This chapter explains unpacking, receipt inspection and claims, and the general procedures for QA-ES setup.

## Receipt, Inspection 	and Return

1.	Inspect the outer box for damage.

2.	Carefully unpack all items from the box and check to see that you have the following items:

- QA-ES Electrosurgical Analyzer (PN 14010)

- E-Input Measuring Cable -- Black (PN 11411-B)

- E-Input Measuring Cable -- Red (PN 11411-R)

- Alligator Clamp -- Black (PN 11412-B)

- Alligator Clamp -- Red (PN 11412-R)

- Power Cord (No PN)

- QA-ES User and Service Manual (PN 14025)

3.	If you note physical damage, or if the unit fails to function according to specification, inform the supplier immediately. When METRON AS or the company’s representative, is informed, measures will be taken to either repair the unit or dispatch a replacement. The customer will not have to wait for a claim to be investigated by the supplier. The customer should place a new purchase order to ensure delivery.

4.	When returning an instrument to METRON AS, or the company representative, fill out the address label, describe what is wrong with the instrument, and provide the model and serial numbers. If possible, use the original packaging material for return shipping. Otherwise, repack the unit using:

- a reinforced cardboard box, strong enough to carry the weight of the unit.

- at least 5 cm of shock-absorbing material around the unit.

- nonabrasive dust-free material for the other parts.

Repack the unit in a manner to ensure that it cannot shift in the box during shipment.

METRON’s product warranty is on page ii of this manual. The warranty does not cover freight charges. C.O.D. will not be accepted without authorization from METRON A.S or its representative.

## Setup

1.	Equipment connection is as shown in the typical setup below (for Power Distribution Test).

2.	If you are using an oscilloscope, attach the BNC cable to the Scope Output connector, located on the front of the QA-ES.

3.	If PRO-Soft QA-ES is being used, attach an RS-232 (null modem/data transfer configured) cable to the 9-pin D-sub outlet port located at the rear of the QA-ES. Do not attach the printer cable to the QA-ES.  See below.  However, if you are not using PRO-Soft QA-ES, and are sending directly to a printer for printouts, attach the printer cable to the 25-pin outlet port.

## PRO-Soft QA-ES

PRO-Soft QA-ES is a front-end test automation and presentation tool for METRON's ESU Performance Analyzer.  It allows you to conduct the same tests, but by remote control via an IBM-compatible PC/XT with MS Windows (Version 3.1 or later). Additionally, the program has features to enhance your QA-ES’s performance.

Each of the QA-ES tests can be run independently from PRO-Soft in the “Manual” test mode.  Results are shown on the PC screen during testing, and the user is prompted to set the tested equipment accordingly. At the conclusion of tests, the user may print a report, store the test and results on disk, or both. Combinations of  tests can be created and stored as “Test Sequences.” The program maintains a library of these sequences. In this way you can store and retrieve sequences that are appropriate for each ESU being tested at your facility.

Sequences can then be used independently, or can be attached to a checklist, written procedure, and equipment data in the form of a test “Protocol.”  The equipment data can be entered manually into the protocol, or it may be retrieved by PRO-Soft from a database program, or other equipment files. Protocols can be created easily for each ESU in your inventory, and stored for use. Test protocols with results can be printed, or stored on disk, and the results of testing can be sent back to the equipment database to close a work order and update the service history.

This page intentionally left blank

# Operating QA-ES

This chapter explains the operating controls, switches and menus of the QA-ES, and details how to use them in ESU testing.

## Control Switches		and Connections

Front Panel

| 1. Power Switch | Turns the power on and off. |
| --- | --- |
| 2. Encoder | Sets values according to the specified range and choose between different operations/ measurement ranges. |
| 3. Enter | Admits newly specified information. |
| 4. Cancel | Cancels a new value and returns to previously chosen value. |
| 5. LCD Display | Shows messages, test results and function menus. |
| 6. Function Keys | Fl - F5 are used to select the functions shown on the bottom line of the LCD display, i.e., for selecting the function that is directly above the key. |
| 7. RF-Detect | Indicates when the ESU is activated. |
| 8. Remote | Indicates that REMOTE CONTR. (F4) has been pressed. |
| 9. Scope Output Connector | BNC-cable connector for attenuator signal in real time. (When oscilloscope output is desired.) |

Right Side Panel

| 10. Terminal RED and BLACK | Connection for the electrode outputs of the VAR. LOAD ESU. Active electrode to the red terminal, and neutral electrode to the black terminal. |
| --- | --- |
| 11. Terminal BLUE and BLUE | Additional fixed load resistance of 200 ohms FIXED LOAD 500 watt for serial connection during leakage test. |
| 12. Terminal GREEN and GREEN | Foot switch output switch can be used to trigger the ESU. |

Rear Panel

| 13. RS-232 Serial Port | 9-pin D-sub |
| --- | --- |
| 14. Printer Outlet Port | 14-25 pin D-sub |
| 15. Main Connector | 3-pin SCHUCO-plug |
| 16. Voltage Selector | 115 VAC/230 VAC |
| 17. Fuses | T 200 mA @230 VAC / T 400 mA @ 11 5 VAC |

## QA-ES Menu and 	Function Keys

The QA-ES uses a display, programmable function keys and a setting regulator to provide flexibility and control over the operations. The upper part of the screen displays messages, status and results. The menu bar is at the bottom of the display. The function keys are numbered from Fl to F5.  A function is selected by pressing the key located directly under the Menu Item displayed in the menu bar.

## LCD Display Menu/ 	Messages (Overview)

## LCD Display Menu/ 	Messages (Detail)

1. 	Startup Screen.  The following screen will be displayed for 2 seconds after the QA-ES has been  switched on.

2.	Main Menu

a.   First Menu Bar (Page 1)

b.   Second Menu Bar (Page 2)

3.	SHOW CHOICES (F1).  This function is activated when you see an asterisk (*) in the status field under ‘Mode.’ Choose a test function by pressing UP (F2) or DOWN (F3).   (The encoder can also be used for choosing a test function) Press ENTER (F5) to save it under Mode in the STATUS field. Press CANCEL (F4) to undo.

4.	KNOB PARAM. (F2).  With this function, you can choose between ‘Mode,’ ‘Load’ and ‘Delay’ in the STATUS field. (* marks the active item). If you choose ‘Load’, use the encoder to set the load from 10 ohms to 5200 ohms in steps of:

- 25 ohms from 50 ohms to 2500 ohms.

- 100 ohms from 2500 ohms to 5200 ohms.

Save the selected load in ‘Mode’ under the STATUS field by pressing ENTER (F5). Press CANCEL (F4) to undo. If you choose ‘Delay’, use the setting regulator to set the delay from 200 ms to 4000 ms in steps of:

- 50 ms from 200 ms to 1000 ms.

- 100 ms from 1000 ms to 4000 ms.

Save the chosen delay in ‘Delay’ under the STATUS field by pressing ENTER (F5). Press CANCEL (F4) to undo.

5.	START (F3). When you press on START, the test procedure will begin, and the text in the field ‘Oper.’ will change from ‘Ready’ to ‘Measuring’. If the unit is set to the position for a REM test, this text will change from ‘Ready’ to ‘lncr'. res.’ Press STOP (F3) to stop the test procedure.

6.	SETUP (F4). Here you can set the power distribution level for start, stop and step in ohms.

Choose the ‘Start load’ by using KNOB PARAM (F4). (see stars). Use the encoder to set the level. Save the level by pressing ENTER (F5). Press CANCEL (F4) to undo. Go to ‘End load’ and ‘Step Size’ and repeat the same procedure.

Pwr. distr. Start load is the first load to be used during the measurements; it can be set from 10 ohms to 2100 ohms, with steps of 25 ohms starting at 25 ohms onwards.

Pwr. distr. End load is the last load used in the measurements; R can be set from 525 ohms to 5200 ohms, with steps of 25 ohms from 525 to 2500 ohms and step of  100 ohms from 2500 ohms to 5200 ohms

Pwr, distr. Step Size is the load set with steps of 25, 50, 100, and 200 ohms

Press QUIT MENU (F5) to return to the main menu.

7.	PRINT HEADER (F3). Writes a heading for a new test protocol.

8.	REMOTE CONTR. (F4). Enables you to control the QA-ES through a PC. Required software: PRO-Soft QA-ES.

## Printout

Press PRINT HEADER (F3) before printing out a page if you want it to have a new heading. The QA-ES automatically prints out the test results via the printer output after every measurement. See example below.

## Foot Switch Output

A Foot Switch Output is activated by use of relays (K11), and located on the right side of the unit. This is used to trigger the foot switch input on the ESU being tested.

## Main Switch On/Off

The QA-ES has to be turned off for at least 5 seconds before turning it on again to allow the reset circuit to unload.

# ESU Tests with QA-ES

This chapter explains the tests that can be conducted by the QA-ES on an ESU, as well as the features available with the PRO-Soft QA-ES software accessory.

## Power Distribution

This test checks the power provided by the ESU over a range of load resistances. The QA-ES allows you to specify a range of loads, over which you test the ESU output power to see if it is within the specified limits.

Per IEC 601-2-2 the power output cannot be reduced by more than 10W, or 5% of the minimum power output level. Per ANSI/AAMI HF18-1993 the power output must be within 20% of the ESU manufacturer’s specifications.

|  | Load Resistance Range |
| --- | --- |
| Equipment | IEC | ANSI/AAMI |
| Monopolar | 100 - 1000 ohms | 50 - 2000 ohms |
| Bipolar | 10 - 500 ohms | 10 - 1000 ohms |

Test setup for ESU power distribution test:

## HF Current Leakage

This test checks to see whether or not the active and dispersive leakage currents are within acceptable limits. There are four test setups to accomplish this testing.

Per IEC 601-2-2 and ANSI/AAMI HF18-1993 the ESU shall be operated at the maximum output setting in each operating mode. The limits for the acceptable leakage currents depend upon the test configuration.

| Test Configuration | Limits of Acceptable Leakage Current |
| --- | --- |
| Measured on electrodes | The leakage current should not exceed 150 mA |
| Bipolar | The leakage current should not exceed 1% of the maximum bipolar rated power output. |
| Measured at equipment terminals | The leakage current should not exceed 100 mA. |

1.	Grounded HF Equipment:  Measurements of the HF current leakage.  The ESU is grounded. The test load is 200 ohms and the ESU must be operating at maximum power. The current leakage measured directly at the instrument's terminals must not exceed 100 mA.

Test setup in compliance with IEC 601.2.2, sec. 19.101a, test 1, fig. 102 and sec. 19.102.  (Adopted by ANSI/AAMI HF18-1993)

Test setup in compliance with IEC 601.2.2, sec. 19.101a, test 2, fig. 103 and sec. 19.102. (Adopted by ANSI/AAMI HF18-1993)

2.	HF Isolated Equipment:  Measurements of the HF current leakage from the active and neutral electrodes. The test load  is 200 ohm and the ESU must be operating at maximum power. The current leakage measured directly at the instrument's terminals must not exceed 100 mA.

Active electrode test setup in compliance with IEC 601.2.2, sec. 19.101b, fig, 104 and sec. 19.102. (Adopted by ANSI/AAMI HF18-1993)

Neutral electrode test setup in compliance with IEC 601.2.2, sec. 19.101b, fig, 104 and sec. 19.102. (Adopted by ANSI/AAMI HF18-1993)

## REM Alarm

This test ensures that the ESU will sound an alarm if the resistance between the two neutral electrodes exceeds your specified limit. The program directs the OA-ES to gradually increase the resistance. At a certain value, the ESU should sound an alarm. Test setup for ESU REM alarm test.

This page intentionally left blank

# Control and Calibration

This chapter explains the QA-ES maintenance procedures, including testing and calibration.

## Required Test 	Equipment

- ESU, 200W in 75 ohms with 500 kHz

- Signal generator, 3 MHz, harmonics better than -40 dBc with 500 kHz, 0 dBm

- Digital multimeter

- Digital multimeter, HP 34401 or equivalent

- RMS / Peak Voltmeter, RHODE & SCHWARY URE 3 or equivalent

- Oscilloscope, 10 MHz

- Electrosurgical probe for QA-ES

- Computer (PC) with RS-232C interface

- Printer with parallel interface

- Short circuit SMB terminal, female

- Test cables

## Preparation

Before performing testing and calibration, you must dismantle the housing. This is accomplished by removing the side plates fastened with Velcro( straps. Then, remove the bottom plate by loosening the four screws. Afterwards, move the bottom plate of the isolation box by loosening the 16 screws holding it in place. Adjust the voltage switch at the rear of the QA-ES so that is corresponds to the main voltage (115 or 230V).

## Function Testing

- Power Supply.  Connect the multimeter in series on one of the mains supply leads to measure the current consumption. Turn on the QA-ES. Don't activate any functions. Measure the current. Required value:

230±10%:	80mA -10/+20mA

115±10%:	160mA -20/+40mA

- User Interface and Display. Check that the display and the user interface are working normally.

- Cooling Fan.  Check that the fan increases speed when you press on Start (F3) in the main menu and that the speed decreases about 10 seconds after you press STOP (F3).

- Serial and Parallel Interface. QA-ES can be connected to a printer and a PC. Check that booth the serial and parallel interfaces work.

- Foot Switch.  Connect a multimeter to the foot switch output on the QA-ES (green). Check that the relay for the foot switch is connected when you press START (F3).

- Fixed Load.  Measure the value of the fixed load with a multimeter (blue).  Required value: 200 ohms ±10 ohms.

- Housing Isolation.  Use a multimeter to check the isolation between the housing and the measuring inputs.  It is important to check all of the terminals on the housing: ‘Var. Load', 'Fixed Load' and 'Foot Switch’.

## Calibration

- Offset Voltage.  Short-circuit the RF input on the sampling unit (J4). All the measurements and reference values in this section refer to the sampling unit. Measure the voltage at each of the following points and adjust them to the following values with the potentiometers specified between brackets. The voltage should be set as close to 0V as possible.

| TP1 - GND (no adjust) | TP3 - TP2 (R3) | TP4 - GND (R2) | TP5 - GND (R1) | TP6 - GND (R4) |
| --- | --- | --- | --- | --- |
| OmV (100 (V | O (V (50 (V | O (V (100 (V | O (V (500 (V | OmV (5 mV |

- Load Resistance.  Set the QA-ES in calibration mode by holding the function key F1 while switching the instrument ON. Continue to press on the key until the main menu appears. Connect a multimeter (hp 34401A) to the 'Var. Load' input.  The resistances are calibrated by turning the small wheel until the display shows a value as close as possible to the value read. To calibrate the next resistance, press on UP (F1) or DOWN (F2).

- Measuring Resistance.  Set the QA-ES in calibration mode.  Before mounting it, the measuring resistance R1 on the Load Board should be measured with hp 34401A.  The measured value should be noted near R1 on the Load Board. Use this value for R1 in the following equation and for ‘Meas. Reas.’ in the calibration menu by using the wheel. Press QUIT (F5) to save calibrated values.

- Measuring Device

Measuring The Effect.  Set the QA-ES in test mode by holding F2 while switching the power on. Connect the signal generator to the RF input on the sampling unit. The signal applied is a pure sine of 500 kHz. URE 3 is connected in parallel with the sampling unit to measure the applied signal. Press START (F3) and read the measured effect in dB. Vary the level of the applied signal from -16 dBm to +10 dBm and find the average value for the difference between the value read on the QA-ES and the applied value measured with URE 3. Use the term below to calculate a linear percentage value from the difference value in dB.

KdB:  Average value for the difference between the value read  on QA-ES and the applied value measured with URE 3.

Clin:  Linear value of kdB.

The following term will lead to a new value for Q_gain.

Turn the wheel and press ENTER to specify a new value for Q_gain. Repeat the test until the average value between the QA-ES and URE 3 is smaller than 0.1 dB. This corresponds to an error of 2.3%.

Peak Detector.  Once the measuring device used in measuring the effect has been calibrated, you can calibrate the peak detector. Use the same test setup as for calibrating the effect. Measure with a signal level of 200 mVp and 1 VP applied from the generator. The applied amplitude is measured with URE 3. The difference between the value read on the QA-ES and URE 3 should not exceed 10% otherwise the P_offset and N_offset in the calibration menu must be changed until this requirement is met. If the positive peak value read is too small P_offset should be reduced. N_offset should be reduced if the negative peak value is too low. Press QUIT (F5) to save calibrated values.

- Measuring Accuracy of the total system.  Connect the coaxial cable from the Load Board to the RF-input (J4) on the Sampling Unit. Put the ground plate of the isolating box back into place. Connect a high voltage probe in the port for variable load resistance on the QA-ES. The probe's coaxial cable should be connected to URE 3. The banana adapters of the probe are connected with the high voltage generator. The following mathematical formula presents the applied effect in dBm as a curve of the effect on URE 3.

R1: 	Resistance in the high voltage probe.  Specified on the probe. Nominal value 47 kohms.

R2:	Measuring resistance in the high voltage probe. Specified   on the probe. Nominal value 37.5 ohms.

The measuring accuracy of the QA-ES should be checked for the following resistance values and the corresponding signals applied.

| R | Applied signal measured on URE3 |
| --- | --- |
| [ohms] | [W] | [dBm] |
| 75 | 200 ( 50W | 52.0 - 54.0 |
| 300 | 150 ( 50 W | 50.0 - 53.0 |
| 1500 | 100 ( 50 W | 47.0 - 52.0 |

Calculate the average value for the difference between QA-ES and URE 3 and correct the value for Q_gain in the calibration menu by following the instructions given before under item 4 in this section.

Requirement:

In linear form this corresponds to:

- Scope Output.  Connect the high voltage generator to 'Var. Load' and the oscilloscope to 'Scope Output'. Set the QA-ES to continuous mode and start measuring by pressing on Start. Activate the high voltage generator and measure the peak-to-peak voltage on the oscilloscope. Check that this value equals the current flow read on the QA-ES.

Nominal value:	4.5 ( 0.4 V/A

# Component Functions and Parts

This chapter provides a detailed description of the functions of the main components of the QA-ES, as well as a parts list for cross-reference. Reference is made to the component location and schematic diagrams to assist servicing personnel. These diagrams are foldouts, and are located in Appendix B.

## Processor Board

(Refer to QA-ES Processor Board Component Location Diagram and Schematic Diagrams 1 and 2) The Processor Board is installed on the inside of the front plate. It comprises the:

- power supply;

- microprocessor system;

- display;

- function keys;

- interface towards Sampling Unit, and;

- RS 232 port and printer port

The Processor Board activates and controls the sampling procedure in the QA-ES. The data is routed back from the Sampling Unit to the microprocessor, where the results will be calculated before being displayed.

1.	Power Supply

(See QA-ES Processor Board Schematic Diagram 2). The Processor Board receives 12 VAC from the Transformer located at the rear of the QA-ES via J4. The voltage is then converted with diode D9 and filtered through C1. Voltage regulator U8 supplies the circuits with +5 V. The voltage is adjusted with resistors Rl0 and R11. Schottky diode D4 protects the 5V power supply to the printer interface to avoid any power flow between the printer and the Processor Board while the QA-ES is switched off. Capacitive switch regulator U9 generates -12 V from +12 V. V-comparator U10 monitors the +5V voltage and sends a reset signal to the microprocessor when the voltage drops under 4.75 V.

Transistor Q4 controls the voltage to the compressor at the rear of the QA-ES. The compressor is connected to J5.  The microprocessor issues a digital signal at PWMA, depending on the ventilation speed desired. Transistor Q2 controls the basic power in Q4. When Q4 is turned off (low ventilation speed), the compressor receives power via transistors R8 and R17, leading to a voltage drop.

2.	Microprocessor System

(See QA-ES Processor Board Schematic Diagram 1). The microprocessor system is in PCMCIA card format (U1). It comprises a MC68HC16 microprocessor with I/O, 256K Flash ROM and 64K RAM.  It can be reprogrammed or replaced when upgrading the software. The processor's I/O includes an asynchronous serial port (ACIA), synchronize serial port, parallel I/O, A/D converter and PWM output. In addition, 8 bit of the data bus are accessible, as well as chip-select lines for direct access to the external I/O.

The processor is timed with a 16.67 MHz timer frequency, controlled by an internal crystal in the component. To check the timer frequency, a 1024 Hz square signal is applied to one of the PWM outputs.  The signal can be measured at pin TP 1.

Ul3 is an EEPROM connected to the processor component via a parallel I/O. Ul3 saves the calibration parameters for the Sampling Unit. The parameters can be stored independently of the processor component so that the QA-ES does not have to be re-calibrated when upgrading the software.

3.	Function Keys

(See QA-ES Processor Board Schematic Diagram 1). The QA-ES is operated by touch keys and a universal digital encoder. The encoder, or 'knob', will be given different functions depending upon the parameters that are to be changed.  The encoder is connected to two touch keys with unchangeable functions, 'ENTER' and 'CANCEL'. The QA-ES is operated from 5 'soft-keys' linked to menus shown on the display.

Latch U2 buffers data from touch keys SW1-SW7. If you press one of the keys, power will flow into the base at Q1 via RP1, thus controlling an interrupter input on the processor unit. When the interrupter input is activated the processor reads latch U2 latch to find out which key was pressed. The same interrupter input is activated via diodes D1 and D2 when operating the encoder. The encoder is read via I/O ports on the processor component.

4.	Display

(See QA-ES Processor Board Schematic Diagram 1). The QA-ES is equipped with a 240 x 64 dot graphic display with a built-in character generator.  When in character mode, the display shows 40 x 8 characters. The display is controlled by the processor unit via the data bus. The display's contrasting voltage is regulated with voltage regulator U12. Voltage (and contrast) are controlled with the potentiometer R15. This is the only point of adjustment at the processor board.

The display offers EL background light to make it easier for the user to read. U11 is an oscillator generating an operating voltage of about 90 VAC for the EL component.

5.	Interface Towards Sampling Unit

(See QA-ES Processor Board Schematic Diagram 1). All data transferred between the Sampling Unit and the Processor Unit is in serial format to simplify the optical interface on the sampling board. Communication is controlled by the synchronous serial interface on the processor unit. The signals are transmitted via pin board base J1.

6.	Serial Port

(See QA-ES Processor Board Schematic Diagram 2). The serial port is adapted to a 9-pin RS-232C format. The port is set to 9600 baud, 8 data bits, 1 stop bit and no parity. RS-232 driver U6 drives the data signals. The handshake is software-oriented. The command responses are returned via the D-sub terminal.

7.	Printer Output

(See QA-ES Processor Board Schematic Diagram 2). The Processor Board’s printer output has a standard 25-pin D-sub contact for Centronix interface. The output is built around 3 HCMOS circuits; U3, U4 and U5. The circuits are connected to the data bus and I/O ports of the processor unit. U3 is a latch for the 8 parallel data lines. U4 is the driver for the outgoing commands, while U5 acts as a buffer for incoming commands. RP3 comprises pull-up resistances for the input lines. All signals to the printer output are filtered to reduce high frequency radiation.

## Sampling Unit

The QA-ES Sampling Unit is placed inside the internal protection box of the instrument. The card is fixed vertically on the right side of the protection box. The unit comprises:

- power supply;

- measuring device, and;

- interfaces towards the Processor Board and the Load Board.

The Sampling Unit converts the applied RF signal to a low frequency signal proportional to the mean square of the RF signal. The peak value of the applied signal is also measured. The measuring values are sampled with a 12 bits A/D converter and the data is transmitted to the Processor Board for calculation and presentation on the display.

1.	Power Supply

(See QA-ES Processor Board Schematic Diagram 1). The Sampling Unit receives -10 VDC and +10 VDC from the Load Board via J1). Voltage regulators U17 and U19 provide, respectively, +6V and -6V for the RMS DC converter. Voltage regulators U18 and U20 provide the peak detector with +9V and -9V. Voltage regulator U16 supplies the opto-coupler in the interface towards the Processor Board with + 5V.

2.	Measuring Device

The measuring device consists of two blocks:

- a RMS DC converter for current measuring, and;

- a peak detector for measuring the peak voltage in the signal.

The peak detector is divided into two actions for measuring the positive and negative peak values.

3.	Input Filtration

At the measuring system input, the incoming signal is filtered down to a lower level through a filter based on L1, L2, C5, C6 and C7. The 3dB frequency is set to 10 MHz. The filter can be found on diagram 1 of 3.

4.	RMS DC Converter

(See QA-ES Processor Board Schematic Diagram 1). The RMS DC converter is based upon an analog multiplicator from analog devices AD834, U15. U15 and U25 form a circuit that calculates the mean square of the applied signal. A filter, consisting of C8, C9. C50, C51, R35 and R36, determines the constant time factor for the integration, which will correspond to the amplitude of the outgoing signal at U25. Further calculations of RMS values and current are carried out in the software on the Processor Board.

Two amplification steps following the mean square circuit ensure optimum dynamics in the measuring system. Both steps U13 and U14 provide 20 dB amplification.  By measuring the signal before U13, and after U13 and U14, you can determine which signal level makes the best use of the A/D converter dynamics.

5.	Peak Detector

(See QA-ES Processor Board Schematic Diagram 3). The peak detector is divided into two sections: one for detecting the positive peak voltage, and the other for detecting the negative peak voltage.  The detector is based on a transistor diode connection.  The positive peak voltage is detected by transistors Q1 and Q4, to which are attached various components. The same goes for the negative peak voltage with Q2 and Q5.  The positive peak voltage is built up over the C27 capacitor, whereas the negative peak voltage is built up over the C28 capacitor.  U23 drains small amounts of current from C27, and provides the AID converter with correct polarization and low impedance for the positive peak detector.  U22 does the same for the negative peak detector. The peak detector ran be reset by short-circuiting C27 and C28 via octol analog switch U24 (See QA-ES Processor Board Schematic Diagram 2).

6.	A/D Converter

The measuring device uses a 12-bit A/D converter from linear, Ul. This circuit measures both RMS values, and positive and negative peak voltage. Each of the signals is multiplexed into the A/D converter via the U24 switch.

The sampling speed depends on the Processor Board's reading speed. Schottky diode D7 protects the A/D converter from incoming negative signals.

7.	Interface Towards Load Board and Processor Board

The measuring device and load resistances in the QA-ES are isolated from the processor unit and the housing by a galvanic shield. This protects the user if a fault in the QA-ES produces ground currents during the measurements. This interface is based upon opto-couplers and various mains transformers for the user interface and the measuring device in the QA-ES.

Data is transferred digitally between the Sampling Unit and the Processor Board in serial form via J3 to simplify the optical interface. Opto-couplers U5, U6, U7 and U8 are used in transmitting signals from the processor system to the Sampling Unit. The Opto-coupler U9 transfers the measurement from the Sampling Unit back to the processor system on the Processor Board.

The relays used in choosing load values on the Load Board are also controlled via terminal J3 on the Sampling Unit. The commands are then transmitted from the Sampling Unit to the Load Board via terminal J2.

8.	High Voltage Protection

Diodes D8-D15 are a protection device against high voltage signals entering the Sampling Unit.

## Load Board

(Refer to QA-ES Load Board Component and Schematic Diagrams). The Load Board is located inside the internal protection box in the QA-ES. The card is fixed vertically inside the box.  The unit includes a power supply, load resistances with corresponding control relays, and interface towards the processor system Sampling Unit, a measurement resistor with attenuators and measuring Transformer for the scope output.

The Load Board forms the load for the ESU being tested. The load can vary from 10 ohms to 5200 ohms, or be fixed to 200 ohms.  A high voltage relay inside the Load Board (K1) connects the load to or from the ESU.

1.	Power Supply

The Load Board is supplied with 2 x 9 VAC from the Trafo in the front part of the protection box via J14.  The voltage is converted with diode D9, and filtered through Cl and C2. U1 is a voltage regulator, providing the relay drivers +5 V.  U1 also supplies Sampling Unit opto-couplers U10, U11 and U12 with 5 V. Cl and C2 provide relay drivers U2 and U3 with +10 VDC and -10 VDC. These voltages are used for controlling the relays on the Load Board.

2.	Load Resistances

There are two load resistances on the Load Board, seen from the ESU being tested. The first has a fixed load of 200 ohms, and is used to measure the current leakage. The other is a variable load resistance, which varies in steps of 25 ohms, from 10 ohms to 5200 ohms.

3.	Fixed Load Resistance

The fixed load resistance is based on the two 100 ohms non-inductive resistances (R17 and R18) connected in series. The total load can be 350 W continuously.

4.	Variable Load Resistance

The variable load resistance is also based on a non-inductive resistance (R3 - R16). In addition, relays (K2 - K10) have been used to offer the choice between different combinations of resistances. You can therefore obtain any value between 10 and 5200 ohms.

5.	Measuring Resistance and Attenuator

The measuring signal transmitted to the Sampling Unit is drained over a 2.0 ohms resistance (R1) in series with the variable load resistance (R3 - R16). Before the signal is transmitted to the measuring device, it passes through a variable attenuator, based on a 10 dB attenuator (R20 - R22) and a 20 dB attenuator (R26 - R30). By connecting these attenuators in series, you can obtain an attenuation of 30 dB.

6.	Interface Towards Sampling Unit

When the signals pass through an optical barrier between the processor system and the rest of the electronics, the Sampling Unit's F1 terminal controls the relays, choosing the load and attenuator values on the Load Board.

7.	Scope Output

The scope output is based on a measuring Transformer L1, mounted on the Load Board. The measuring Transformer is terminated with a 50 ohms resistance R31.

8.	Foot Switch

The output foot switch is based on a relay (K11), and is used in triggering the ESU being tested.

## Component Parts

| COMPONENT PART | TYPE/VALUE | QTY. | DIAGRAM REFERENCE |
| --- | --- | --- | --- |
| HOUSING: |  | 1 |  |
| Transformer ring core 1x12V 1A | ULVECO AA81002 | 1 |  |
| Transformer E-core 10kV 2x9V | ELTRAFO | 1 |  |
| D-Sub 25p female | ELFA 43-674-54 | 1 |  |
| D-Sub 9p male | ELFA43-673-97 | 1 |  |
| Flat cable contact 10-polt | ELFA 43-646-00 | 1 |  |
| Flat cable contact 16-polt | ELFA 43-646-26 | 2 |  |
| Flat cable contact 26-polt | ELFA 43-646-42 | 1 |  |
| BNC Straight bulkhead jack | R141306000 | 1 |  |
| SMB coax. conn. Straight plug | Rl14082000 | 1 |  |
| Security contact Red | Flat 4.8mm JHSupp 404-171 | 1 |  |
| Security contact Black | Flat 4.8mm JHSupp 404-137 | 1 |  |
| Security contact White | Flat 4.8mm | 2 |  |
| Security contact Green | Flat 4.8mm | 2 |  |
| Ventilator 12 VDC 92mm | Panasonic FBA09A12H1A | 1 |  |
| Protection grate for ventilator | Sunon FG-9 | 1 |  |
| Turning knob |  | 1 |  |
| Voltage regulator | C&K V802-12-SS-05-Q | 1 |  |
| Apparatus input with netfilter | Corcom lED4 | 1 |  |
| Safety fuse 5x20 mm | Schurter 0031.1081 FEF | 2 |  |
| Safety fuse 5x20 mm | Schurter 0031.1363 FIO | 2 |  |
| Net switch | C&K DM22-J1-2-S2-05-N-Q | 1 |  |
| Cable shoe |  | 13 |  |
| ABIKO Ring cable shoe |  | 1 |  |
| Screw dimension M3 |  |  |  |
| Flat cable |  | 1 |  |
| Flat cable to printer | 25-polt 44 cm | 1 |  |
| Flat cable to serial port | 9-polt 44 cm | 1 |  |
| Flat to measuring board | 16-polt 33 cm | 1 |  |
| Coax. cable IRG174 | 45 cm Farn 125326 | 1 |  |
| Countersunk flat headed screw | DIN 965 M4x12 | 14 |  |
| Countersunk flat headed screw | DIN 965 M3x6 | 8 |  |
| Screw slot SH | DIN 84A M3x10 | 4 |  |
| Screw slot SH | DIN 84A M3x8 | 12 |  |
| Screw slot SH | DIN 84A M4x10 | 4 |  |
| Screw slot SH | DIN 84A M4x35 | 1 |  |
| Screw recessed head poz | DIN 7985 M2-5x10 | 8 |  |
| Screw recessed head poz | DIN 7985 M3x6 | 31 |  |
| Screw recessed head poz | DIN 7985 M3x8 | 4 |  |
| Screw recessed head poz | DIN 7985 M3x10 | 4 |  |
| Screw recessed head poz | DIN 7985 M4x30 | 4 |  |
| Nut M3 |  | 6 |  |
| Nut M4 |  | 4 |  |
| Safety nut M4 |  | 1 |  |
| Spring washer M4 |  | 4 |  |
| Washer M4 | 12mmø DIN 9021 | 1 |  |
| Spacer for fixing the front plate | 5 |  |
| Contact plate |  | 1 |  |
| RF protection box |  | 1 |  |
| Front plate w/5 screws for circuit board |  | 1 |  |
| Rear plate |  | 1 |  |
| Front foilv |  | 1 |  |
| PROCESSOR BOARD: |  |  |  |
| Printed circuit board | Elprint AR075 | 1 |  |
| Micro Module | 16MM16 | 1 | U1 |
| Latch | 74HC574N | 3 | U2, U3, U5 |
| Port | 74HC05N | 1 | U4 |
| RS 232-driver | MC145406P | 1 | U6 |
| Voltage Reg. | LM337LZ | 1 | U12 |
| Voltage Reg. | LM317T | 1 | U8 |
| V-converter | ICL7662CPA | 1 | U9 |
| V-comparator | Mc34064P-5 | 1 | U10 |
| V-converter | NeI-D32-49 | 1 | U11 |
| EEPROM | X24C02P | 1 | U13 |
| Transistor | BC547B | 2 | Q1Q2 |
| Transistor | BD140 | 1 | Q4 |
| Diode | 1N4148 | 4 | D1, D2, D5, D10 |
| Diode | 1N4002 | 1 | D3 |
| Schottky diode | 1N5819 | 1 | D4 |
| Bridge | 4A 200v | 1 | D9 |
| LCD-display | Optrex DMF 5005N-EW | 1 |  |
| Encoder | Bourns ECWIJ-B24-BC0024 | 1 | SW8 |
| LED | Gul 3mm Sie LY3360-K | 1 | D7 |
| LED | Rød 3mm Sie LR 3360-FJ | 1 | D8 |
| Resistor | 22R 1% 0.5W | 2 | R8, R17 |
| Resistor | 240R 1% 0.5W | 1 | Rll |
| Resistor | 330R 1% 0.5W | 2 | R4, R5 |
| Resistor | 750R 1% 0.5W | 1 | R104 |
| Resistor | 1K0 1% 0.5W | 1 | R9 |
| Resistor | 2K2 1% 0.5W | 1 | R14 |
| Resistor | 6K8 1% 0.5W | 1 | R16 |
| Resistor | 10K 1% 0.5W | 5 | R2, R7, R12, R18, R19 |
| Resistor | 22K 1% 0.5W | 2 | R3R13 |
| Resistor | 100K 1% 0.5W | 1 | R1 |
| Resistor pack | 5x2K2 SIL | 1 | RP4 |
| Resistor pack | 8x4K7 SIL | 1 | RP2 |
| Resistor pack | 8x10K SIL | 1 | RP3 |
| Resistor pack | 8x47K SIL | 1 | RP1 |
| Varistor | 20V | 1 | R6 |
| Trimpot | 10K 1-tørn Cermet | 1 | R15 |
| Multilayercond. | 100nF 50V | 6 | C8 |
| El.lytt cond. | 10 F 25V | 4 | C2, C3, C5, C6 |
| El.lytt cond. | 220 F 25V | 1 | C7 |
| El.lytt cond. | 2200 F 25V | 1 | C1 |
| Tantal cond. | 1 F 35V | 1 | C4 |
| EMI-Filter | Murata DSS306-91YSS102M1 | 19 | FI1 - FI19 |
| PCM-ClA slot |  | 1 | (U1) |
| Display-connection | 4 pin Berg 71991-410 | 1 | (DISP1) |
| Display-connection | 20 pin Berg 71991-410 | 1 | (DISP1) |
| Board pin base | 16 pin | 1 | J1 |
| Board pin base | 10 pin | 1 | J2 |
| Board pin base | 26 pin | 1 | J3 |
| Pin base for display |  |  | (DISP1) |
| Pin base | 36 pol | 5 | TP1-TP7 |
| Screw clip basic part | 2 pol | 2 | J4J5 |
| Screw clip | 2 pol | 2 | (J4J5) |
| Touch keys |  | 7 | SW1 - SW7 *1 |
| Cover for touch keys |  | 7 | (SW1-SW7) |
| Safety holder | Schurter 0031-8201 OGN | 1 | F1 |
| Safety fuse | 600mA Slow | 1 | (F1) |
| Heat sink | Elfa 75-612-44 | 1 | (Q4) |
| Screws | M2 x 10 | 2 |  |
| Nuts | M2 | 2 |  |
| Nylon screws | M3 x 12 | 4 |  |
| Nylon nuts | M3 | 12 |  |
| SAMPLING UNIT: |  |  |  |
| Printed circuit board | Elprint AR-069 | 1 |  |
| Op.amp High Current | AD844AN | 1 | U21 |
| Op.amp High Speed | AD711JN | 1 | U25 |
| Op.amp Railto Rail | AD820AN | 2 | U13, U14 |
| Op.amp Dual | LMC662CN | 2 | U22, U23 |
| Analog Multiplier | AD834JN | 1 | U15 |
| AiD-converter 12-Bit 250 kHz | LTC 1272- 5CCN | 1 | U1 |
| Octal Analog Switch Array | DG485DJ | 1 | U24 |
| Shift Register | 74HC589AN | 2 | U2, U3 |
| Quad Nand port | 74HC00AN | 1 | U4 |
| Opto Coupler | CNW2611 | 8 | U5- U12 |
| Voltage Regulator | LM317T | 1 | U16 |
| Voltage Regulator | LM317LZ | 1 | U17 |
| Voltage Regulator | LM2941T | 1 | U18 |
| Voltage Regulator | LM337LZ | 1 | U19 |
| Voltage Regulator | LM2991T | 1 | U20 |
| Transistor | 547BNPN | 2 | Q1, Q2 |
| Transistor | 560BPNP | 2 | Q4, Q5 |
| Diode | LL4148 | 12 | D1 - D6, D8- D13 |
| Zener diode | 1N5335. 39V | 2 | D14, D15 |
| Schottky diode | PRLL5819 | 1 | D7 |
| Crystal | 25 MHz | 1 | Y1 |
| Resistor | 4R7 1% 1/8W | 2 | R56, R57 |
| Resistor | 10R 1% 1/8W | 4 | R27, R34, R37, R58 |
| Resistor | 24R9 1% 1/8W | 1 | R24 |
| Resistor | 51R 1% 1/8W | 2 | R61, R72 |
| Resistor | 75R 1% 1/8W | 2 | R25, R26 |
| Resistor | 100R 0.25% 1/8W | 4 | R22, R23, R35, R36 |
| Resistor | 160R 1% 1/8W | 1 | R28 |
| Resistor | 249R 1% 1/8W | 4 | R42, R44, R46, R47 |
| Resistor | 390R 1% 1/8W | 1 | R90 |
| Resistor | 470R 1% 1/8W | 18 | R6 - R21, R64, R67 |
| Resistor | 750R 1% 1/8W | 3 | R5, R41, R55 |
| Resistor | 953R 1% 1/8W | 2 | R43, R45 |
| Resistor | 1K0 1% 1/8W | 3 | R54, R82, R83 |
| Resistor | 1K2 1% 1/8W | 1 | R75 |
| Resistor | 1K5 1% 1/8W | 4 | R48, R65, R73, R81 |
| Resistor | 1K8 1% 1/8W | 4 | R40, R52, R60, R71 |
| Resistor | 2K0 0.25% 1/8W | 3 | R38, R50, R77 |
| Resistor | 2K2 1% 1/8W | 1 | R89 |
| Resistor | 2K7 1% 1/8W | 2 | R63, R68 |
| Resistor | 3K3 1% 1/8W | 1 | R53 |
| Resistor | 3K6 1% 1/8W | 1 | R76 |
| Resistor | 4K7 1% 1/8W | 11 | R49, R59, R62, R69, R70, R78 - R80, R86 - R88 |
| Resistor | 10K 0.25% 1/8W | 2 | R30, R31 |
| Resistor | 18K 0.25% 1/8W | 2 | R39, R51 |
| Resistor | 22K 0.25% 1/8W | 1 | R29 |
| Resistor | 47K 0.25% 1/8W | 2 | R32, R85 |
| Resistor | 10M 1% 1/8W | 2 | R66, R74 |
| Trimpot Multi. | 10K | 4 | R1, R2, R3, R4 |
| Chip ind. | 10% | 2 | L1, L2 |
| (Chip ind. | 2% | 2) |  |
| Cond Multil. lpF0 5% 63V |  | 1 | C22 |
| Cond Multi. | 47pF 5% 63V | 2 | C2, C3 |
| Cond Multi. | 100pF 5% 63V | 3 | C11, C46, C47 |
| Cond Multi. | 150pF 5% 63V | 2 | C5, C7 |
| Cond Multi. | 470pF 5% 63V | 1 | C6 |
| Cond Multi. | lnF0 5% 63V | 2 | C27, C28 |
| Cond Multi. | 10nF 10% 63V | 20 | C10, C11, C23-C26, C29, C31 - C34, C36, C37 - C39,C41, C42, C44, C45, C60 |
| Cond Multi. | 100nF 10% 63V | 8 | C4, C8, C9, C12, C2, C35, C71, C72 |
| Tantal cond. | 1,0 F 20% 15V | 20 | C30, C48, C49, C52, C53, C56 - C59 C61 - C70 |
| Tantal cond. | 10 F 20% 15V | 5 | C14, C16, C17, C18- C20 |
| El.lytt cond. | 10 F 20% 50V | 4 | C13, C15, C54, C55 |
| Condenser | NC |  | C40, C43, C50, C51 |
| Printer board contact | 6-polt hunn | 2 | J1, J2 |
| SMB coax conn. | Right angle PCB recept. | 1 | J4 |
| Stifflist kort | 16 pin | 1 | J3 |
| 18/36 pin base | 36 pol (testpoints split) |  | TP1 -TP18 |
| Screw | M3x6 Elzink Pan | 3 |  |
| Nut | M3 Elzink | 3 |  |
| LOAD BOARD: |  |  |  |
| Printed circuit board | Elprint AR-061B1 | 1 |  |
| Voltage Regulator | LM317T | 1 | U1 |
| Resistance bridge | 4A 200V GBU4D | 1 | D2 |
| Resistor | 3K0 5% 50W | 1 | R16 |
| Resistor | 9R1 5% 100W | 1 | R3 |
| Resistor | 15R 5% 100W | 1 | R4 |
| Resistor | 620R 5% 100W | 1 | R15 |
| Resistor | 47R 5% 175W | 1 | R5 |
| Resistor | 100R 5% 175W | 4 | R8, R9, R17, R18 |
| Resistor | 200R 5% 175W | 4 | R6, R7, R10, R11 |
| Resistor | 330R 5% 175W | 1 | R12 |
| Resistor | 470R 5% 175W | 1 | R13 |
| Resistor | 1K0 5% 175W | 1 | R14 |
| Resistor | 2R0 1% 10W | 1 | R1 |
| Resistor | 10R 0.1% 0,6W | 1 | R30 |
| Resistor | 26R1 0.1% 0,6W | 2 | R20, R22 |
| Resistor | 34R8 0.1% 0,6W | 1 | R21 |
| Resistor | 82R 0.1% 0,6W | 4 | R26 - R29 |
| Resistor | 100R 0.1% 0,6W | 2 | R19, R25 |
| Resistor | 51R 1% 0,5W | 1 | R31 |
| Resistor | 240R 1% 0,5W | 1 | R23 |
| Resistor | 750R 1% 0,5W | 1 | R24 |
| El.lytt cond. | 10 F 2,5V 25mm rad. | 1 | C3 |
| El.lytt cond. | 2200 F 25V rad. | 2 | C1, C2 |
| Circuit driver | UCN5842A | 2 | U2, U3 |
| High voltage relay | Günther 33911290246 | 1 | K1 |
| Relay 24V | Takamisawa | 10 | K2 - Kll |
| Relay 24V 2-pol | Takamisawa RY-24W-OH-K | 2 | K12, K13 |
| (Relay 24V 2-pol | Takamisawa RY-24W-K | 2) |  |
| Toroid | 3F3 125x75x5 mm ELFA 58-755-21 | 1 | L1 |
| SMB coax conn. | Straight plug crimp | 1 |  |
| SMB coax conn. | Straight PCB recept. | 1 | J13 |
| Printer board contact | 6-polt hannPhoe | 1 | J1, J2 |
| Printer board clip | 3-polt Phoe | 1 | J14 |
| Safety holder | Farn 134477 | 2 | F1, F2 |
| Safety fuse | 600mAT Farn 150202 | 2 | (F1), (F2) |
| High voltage cable | 35 cm 10-/m @ 100 m | 2 |  |
| Coax. cable RG174 | 40 cm Farn 125326 | 1 |  |
| Coax terminator | Double leg right ang. | 1 |  |
| Plastic Spacer | ENISO 3x25 | 38 |  |
| Plastic Spacer | ENISO 3x20 | 6 |  |
| Nylon Screw | M4x07x12 mm | 36 |  |
| Screw pin | DIN 553 M4x22 mm | 14 |  |
| Dekk-kappe todelt |  | 30 |  |
| Bracket for power resistor |  | 30 |  |

This page intentionally left blank

Appendix A - GLOSSARY OF TERMS USED

| Bipolar Electrosurgery | Electrosurgery where current flows between two bipolar electrodes that are positioned around tissue to create a surgical effect. Current passes from one electrode, through the desired tissue, to the other electrode, thus completing the circuit without entering any other part of the patient's body. Neutral plates are not employed in the bipolar technique. Both electrodes are generally of the same size. |
| --- | --- |
| Bipolar Output | An isolated electrosurgical output where current flows between two bipolar electrodes that are positioned around tissue to create a surgical effect in that tissue (usually desiccation). |
| Blend | A waveform that combines features of cut and coag waveforms; current that cuts with varying degrees of hemostasis. |
| Crest Factor | The amount of heat generated is relative to the mean power value. The crest factor depends on the load resistance and is defined as the ratio of peak value to effective value. A sine wave has a crest factor of 1.4 and provides the cleanest form of cutting. |
| Diathermy, also Surgical Diathermy; Electrosurgery | A surgical technique used to cut or coagulate cellular tissue. To avoid muscle contractions, only high frequency currents and voltages of more than 100 kHz are used. The electric current directs the heat into the tissue. The patient is connected to two electrodes, allowing the current to flow through the body. The active electrode will generate a large amount of heat, due to the high current density and the small surface of the electrode. |
| ESU | Electrosurgical Unit. This is a term which is inclusive of both the electrosurgical generator and its connecting cables. |
| Cut (Cut Mode, Pure Cut) | A low voltage, continuous waveform optimized for electrosurgical cutting. |
| Isolated Output | The output of an electrosurgical generator that is not referenced to earth ground. |
| Monopolar Electrosurgery | A type of electrosurgery involving a small (active) electrode and a large neutral (neutral plate) electrode. The small surface of the active electrode provides very good results in coagulating and cutting. The neutral plate of modern units is split, thus controlling the circuit, including the contact between electrodes and patient. |
| Monopolar Output | A grounded or isolated output on an electrosurgical generator that directs current through the patient to a patient return electrode. |
| Resistance (Impedance) | Resistance to the flow of alternating current, including simple direct current resistance and the resistance produced by capacitance or inductance. The resistance of a material is its tendency, measured in ohms, to oppose the flow of electric current or, viewed another way, the material's tendency not to conduct the current. |
| REM | Return Electrode Monitor. |

This page intentionally left blank

Appendix B - DIAGRAMS

Processor Board Component Location Diagram	B-2

Schematic Diagram Part 1 (Processor Board)	B-3

Schematic Diagram Part 2 (Processor Board)	B-4

Sampling Unit Component Location Diagram	B-5

Schematic Diagram Part 3 (Sampling Unit)	B-6

Schematic Diagram Part 4 (Sampling Unit)	B-7

Schematic Diagram Part 5 (Sampling Unit)	B-8

Load Board Component Location Diagram	B-9

Schematic Diagram Part 6 (Load Board)	B-10

| USA 1345 Monroe NW, Suite 255A Grand Rapids, MI 49505 Phone: (+1) 888 863-8766 Fax: (+1) 616 454-3350 E-mail: support.us@metron-biomed.com | FRANCE 30, rue Paul Claudel 91000 Evry, France Phone: (+33) 1 6078 8899 Fax: (+33) 1 6078 6839 E-mail: info@metron.fr | NORWAY Travbaneveien 1 N-7044 Trondheim, Norway Phone: (+47) 7382 8500 Fax: (+47) 7391 7009 E-mail: support@metron.no |
| --- | --- | --- |
| From: (name) |  | Phone: |  |  |
| Address: |  | Fax: |  |  |
|  |  | E-mail: |  |  |
|  |  | Date: |  |  |
| Error Report |
| Product: |  | Version: |  | Serial no.: |  |
| Description of the situation prior to the error: |
| Description of the error: |
| (METRON AS internally) |
| Comments: |
| Received date: | Correction date: | Ref No. | Critical | Normal | Minor |

| USA 1345 Monroe NW, Suite 255A Grand Rapids, MI 49505 Phone: (+1) 888 863-8766 Fax: (+1) 616 454-3350 E-mail: support.us@metron-biomed.com | FRANCE 30, rue Paul Claudel 91000 Evry, France Phone: (+33) 1 6078 8899 Fax: (+33) 1 6078 6839 E-mail: info@metron.fr | NORWAY Travbaneveien 1 N-7044 Trondheim, Norway Phone: (+47) 7382 8500 Fax: (+47) 7391 7009 E-mail: support@metron.no |
| --- | --- | --- |
| From: (name) |  | Phone: |  |  |
| Address: |  | Fax: |  |  |
|  |  | E-mail: |  |  |
|  |  | Date: |  |  |
| Improvement Suggestion |
| Product: |  | Version: |  |
| Description of the suggested improvement: |
| (METRON AS internally) |
| Comments: |
| Received date: | Correction date: | Ref No. | Critical | Normal | Minor |

| USA 1345 Monroe NW, Suite 255A Grand Rapids, MI 49505 Phone: (+1) 888 863-8766 Fax: (+1) 616 454-3350 E-mail: support.us@metron-biomed.com | FRANCE 30, rue Paul Claudel 91000 Evry, France Phone: (+33) 1 6078 8899 Fax: (+33) 1 6078 6839 E-mail: info@metron.fr | NORWAY Travbaneveien 1 N-7044 Trondheim, Norway Phone: (+47) 7382 8500 Fax: (+47) 7391 7009 E-mail: support@metron.no |
| --- | --- | --- |

# Text boxes

|  | CAUTION |
| --- | --- |

The QA-ES uses a 3-wire power cord and plug for the user's safety. Use this power cord in conjunction with a properly grounded electrical outlet to avoid electrical shock.

NOTE

PRO-Soft QA-ES has its own user manual, which contains all the information concerning the program. If you order a demonstration version of the program you also receive the manual.

WARNING!

HIGH VOLTAGES ARE CAPABLE

OF CAUSING DEATH!

USE EXTREME CAUTION WHEN PERFORMING TESTS AND CALIBRATION. USE ONLY INSULATED TOOLS WHEN THE UNIT IS PLUGGED IN, AND THE CASE HOUSING IS OFF.

NOTE

The QA-ES should be switched on for a minimum of 15 minutes before the test starts to ensure stable working temperatures.

NOTE

Measurements should be performed only on main supply, and not both 230V and 115V.

NOTE

It is important to calibrate the resistance of the measuring leads on the multimeter itself before starting the test.

NOTE

Remember to set the reference impedance in URE 3 to the same value as the load resistance used for measuring.

NOTE

Remember that the scope output is not calibrated!