「RF-Sputter Coater Machine」　User Manual

CANON ANELBA ENGINEERING E-200S

December 22, 2009

Overall View	Main Unit	Control Panel

１．Utility Operation

- Turn on the breaker of the switch panel (3 phase, 200V).

- Turn on the compressor and confirm that the air pressure is 0.5 - 0.8 MPa.

- Open the container valve for Argon gas Ar and set the secondary pressure to 0.1 MPa. Note: Opening the container valve by 1/4 rotation is more than enough.

- Fully open the valve of the cooling water.

- Open the two red valves on the left side of the main unit. The cooling water will be supplied to the target and the substrate.

- For the red valves, horizontal position = CLOSE, vertical position = OPEN.

- The green valves are use to bypass the cooling water. These valves are normally CLOSE.

- In order to heat the substrate, set the right red valve to CLOSE position. At the same time, OPEN the right green valve. After that, the recessed water inside the pipe will be blown out by the compress air (~3 minutes).

- Turn ON the main power supply.

- The ALARM will be turned on. Please press the RESET button to cancel the ALARM.

Switch Panel	Compressor	Argon Gas Regulator

２．Operation of Vacuum Pump

- Turn ON the power switch located at the upper part of the control panel. There are two vacuum level sensors, one for low vacuum level and another one for high vacuum level.

- Please confirm that the manual orifice of the coating chamber and Turbo Molecular Pump (TMP) is CLOSED. Rotate the manual orifice to the CLOSE direction lightly during confirmation. DO NOT tighten the manual orifice strongly.

- Press “MONITER” on the touch panel (Don’t ask me why they put “MONITER” instead of “MONITOR”), Then press “Menu” to show the “SELECTION” display.

- Press “PUMP” on the “SELECTION” display to show “PUMPING” display.

- Press “TMP” on the “PUMPING” display to turn ON (green) the exhaust system.

- The exhaust system is operated automatically. First, it operates the Rotary Pump (RP), and then it operates the TMP.

- When TMP is operational, you can hear the “KEEN…” sound.

- At this point, the coating chamber has not reached the vacuum condition yet.

Vacuum gage　　　　　　　Manual orifice (CLOSE)　　　　　　　　　　Touch panel

MONITER display　　　　     SELECTION display	   PUMPING　display	３．Setting Up the Substrate

- Press “Menu” on the touch panel. At the “SELECTION” display, press “TRANS”.

- At the “T/S & ROT CONTROL” display, confirm that PV=1 at “SUBST ROT” and PV=100.0 (mm) at “TS POS”.

- Press “Menu” on the touch panel. At the “SELECTION” display, press “PUMP”. Confirm that “MV” and “RV” is CLOSED. Then, confirm that “SP” is OFF (white).

- TMP will be damaged by the exposure to the air during operation. Therefore, this confirmation is very important.

- Open the VENT valve (Leak valve) “VV” of the coating chamber. The air will flow into the coating chamber.

- After the coating chamber is filled with air (atmospheric pressure), open the lock by rotating the clamp. Then, lift the hatch slowly by holding its support pole.

- Confirm the type of materials in target 1, target 2 (magnetic), and target 3.

- Target 1 and 3 are for non-magnetic materials while target 2 is for magnetic material.

- Mount the specimen on the specimen holder.

- If you want to coat the materials while rotating the stage, place your specimen close to the middle part of the stage.

- Confirm that no dust/dirt/particles are adhered on the O-ring of the flange.

- If there are dust/dirt/particles, clean it using ethanol.

- Close the coating chamber slowly.

- Close the coating chamber using both hands. Be careful not to hurt your hands.

- Close the VENT valve.

- Close the VENT valve slowly.

- Close and lock the hatch’s lock.

SELECTION display	T/S & ROT display	Height of the substrate holder

（TS POS = 100 mm）

PUMPING display	Main Body: front	Hatch’s lock, support pole,

＜Valve＞ void：CLOSE，Filled：OPEN		and VENT valve

Coating chamber OPEN	Target	Specimen holder

４．Evacuation of coating chamber

- Press “MENU” on the touch panel. At the “SELECT” display, press “PUMP”.

- At the "PUMPING" display, confirm that the “TMP” is on, and then turn on the “SP”. The evacuation of coating chamber will be started.

- At first, the sound of the rotary pump will become louder and then gradually decreased.

- Open the manual orifice, then wait until the pressure of the coating chamber become 10-4 Pa. (Approximately 30 minutes).

- Please open the manual ORIFICE slowly.

- During evacuation, when the pressure does not reach 10 Pa due to air leakage, the error message will be shown after 3 minutes.

- In this case, open the coating chamber and check the O-ring. Then, close the chamber again and restart the evacuation process.

MONITER display　　　　　  　　SELECT display        　　　　　PUMPING display

Manual ORIFICE（open）　　　  　　　Vacuum gage

５．Injection of gas(Argon)

  - Turn off the “FILAMENT” on the vacuum gage.

    - If the filament is left on when the gas is injected into the coating chamber, the filament might be damaged.

  - Confirm that the Argon gas pressure on the secondary pressure is 0.1 MPa. Adjust the amount of pressure using the mass flow valve.

  - Open the stop valve “CV”.

  - Press “MENU” on the touch panel. At the “SELECT” display, press “GAS”. At the “GAS CONTROL” display, define the required SV value（Ar: Min 0.2―Max 30sccm.）.

  - After closing the manual ORIFICE, turn on “MFC” then the gas will be injected into the coating chamber.

  - Confirm that the coating chamber pressure is 10 Pa.

  - In addition, it is better to turn on “SP LINK”.

    - During injection, it is possible to change the SV value.

    - If the “SP LINK” is left to be ON, the gas will be automatically stopped after coating process.

  - By adjusting the manual ORIFICE, the pressure inside the coating chamber can be controlled.

    - (If the conditions for coating with Al are [SV: 30 sccm (8Pa), target distance: 80cm, RF power: 150 Watt, coating time: 1 hour], the coating thickness will be 100-120 nm).

Pressure gage　　　　　　　　　Argon gas regulator　　　　　　　　Ar stop valve CV

SELECT display　　　　　　　 GAS CONTROL display　　Manual ORIFICE (close)

６．Plasma generation and coating control

- At “TRANS CONTROL” display, turn on the “HOME POS” switch.

Set the distance between the target to the substrate as 100 mm.

Afterwards, change the distance to the required position → Please refer to the “Rotation and vertical operation method of the specimen holder” on the attached paper.

To rotate the specimen holder, press “STEP” to change it into “CONTINUE”. Then, press”OFF” to change it into “ON”.

To stop the rotation, press “CONTINUE” to change it into “STEP”. The specimen holder will be automatically stopped. Please note that the distance between the target and the specimen can not be changed while rotating.

- Use the knob under the main body to select the target for coating.

- Rotate the knob to “TGT 1”,”TGT 2”, or “TGT 3” to select the target.

- Turn on the “RF Power Generator”.

- Turn the knob on the “RF Power” panel counter clockwise and set the power input to 0 (zero).

- Turn on the “RF Power” knob. Then, when the knob is set to “0” W, the left side display will show “000Wf”. Then, set the knob to “10Wf”.

If the argon gas and the cooling water is not flowing during this step, the error message will be shown. The process will be intercepted by the “interlock”.

- Set the reflected wave “Wf” at the right side display to zero “0” by tuning the “LOAD” (rough adjustment) and “TUNE” (fine adjustment) knobs at the lower part of the main body.

- Slowly increase the power input while tuning the reflected wave to zero “0”. Please be careful not to exceed 10 percent of the RF power input (i.e. if the Wf is 100 Wf, then the Wr should be less than 10Wr). Repeat this step until the required power input is achieved.

- Confirm the generation of plasma from the window at the main body.

(Warning: do not look into the plasma more than a few seconds. Then, do not forget to close the window).

- At the “SP CONTROL” display, set the target opening time ”□h□m□s” prior to the coating process.

- From”TGT1”,”TGT2”,and ”TGT3”; select the required target and then turn it on.

- Turn on the “TIMER” on the touch panel.

(The target shutter will be open, and then the coating process is started. Leave the machine until the coating process is finished. Drink some coffee or something at 7 eleven!!!! DO NOT GO HOME!!!!).

T/S & ROT display        Knob for target selection　　　　　　　　RF adjustment

Reflected wave adjustment　　　　　　 peeking window　　　　　　　　 SP CONTROL display

７．Taking the substrate

- Normally, the argon gas will be stopped after the coating process. If, the argon gas is still flowing, turn off the “MFC” at the “GAS CONTROL” display. After that, close the stop valve CV.

- Open the manual ORIFICE. The coating chamber will be promptly ventilated. Then, close the manual ORIFICE again.

- Turn off the filament of the pressure gage. After that, turn off the “SP” switch at the “PUMPING” display.

- In case of coating with rotation, make sure to stop the rotation at the “T/S & ROT CONTROL” display. Return the specimen holder to the standard position (100 mm).

- Open the VENT valve to ventilate the coating chamber. Open the hatch’s lock and then open the coating chamber by holding its support pole.

- Takes the substrate carefully.

(Warning: when high temperature is used, be careful not to burn your hand).

- Close the coating chamber.

GAS CONTROL display　　　　　　　　   stop valve CV　　　　　　　　manual ORIFICE

Pressure gage　　　　　　　　　　　PUMPING display　　　　　　　　   T/S & ROT display

Coating chamber (front view)　　　　coating chamber (OPEN)

８．Shutting down the equipment

- At the “PUMPING” display, turn on the “SP” switch.

- After turning to TMP, wait for 1-2 minutes in order to make the coating chamber in vacuum condition.

- At the “PUMPING” display, turn off the “SP” switch.

- Turn off the “TMP” switch at the “PUMPING” display (below the “SP” switch).

- Wait for approximately 4 minutes (until the rotary pump and the turbo molecular pump (TMP) are stopped). After that, turn off the pressure gage. Then, turn off the main controller.

- Close the argon gas valve then turn off the compressor. Then, turn off the cooling water. Finally, turn of the breaker of the main switch panel.

PUMPING display　　  pressure gage　　　　　  main controller　　　 　argon gas regulator

Compressor　　　　　　　　　　　　　　Cooling water　  　　　　　     Switch panel

Translators:

    - Willy Kurnia

    - Namatame Yoshiyuki

    - Osawa Hiroki

ANELBA’s contact person:

森本　潤 （もりもと　じゅん）/Morimoto Jun

401-0397 山梨県南都留郡鳴沢村8532-28

TEL: 0555-20-5573（直通）

FAX: 0555-85-3057

E-mail: morimoto.jun@canon-anelvaeng.co.jp

（2009.6.10）

# Text boxes

Electrical Switch

Pressure Control Valve

Container valve

Mass Flow Valve

Control Panel

Main Unit

Argon gas

Compressor

Secondary pressure

Switch Panel

Main Power Supply

ＲＦ Power Supply

Touch Panel

Pressure gage

Coating chamber

Rotary Pump

Cooling water

Compressed air

RESET

Hatch’s lock

Menu

PUMP

TMP

MONITOR

SP

TRANS

VENT valve

Support pole

Specimen holder

Target 1

Target 3

Target 2

(magnetic)

Shutter

Target 1-3

Specimen holder

Target

Substrate

Cooling Water

Compressed Air

FILAMENT OFF

PUMP

SP

TMP

Menu

Pressure inside the coating chamber

Power switch

Breaker

Pressure gage

Mass flow valve

Stop valve CV

GAS

MFC

CLOSE

SP Link

Pressure of the coating chamber

HOMEPOS

Target selection

Power

RFknob

TUNE

LOAD

Target selection

Up: TIMER;

below: opening time

100.0

PV

Holder position

OFF

CV（CLOSE）

OPEN→CLOSE

FILAMENT OFF

SP OFF

OFF

SUBST ROT OFF

OFF

HOMEPOS ON

OFF

Hatch’s lock

VENT valve

Support pole

Specimen holder

SP

TMP

Power（OFF）

POWER

Breaker

ＳＶ

Power（OFF）

Red: Cooling water (OPEN)

Green: Compressed air (CLOSE)

Left: Target; Right: Substrate

Touch Panel