| Procedure | BD MAX™ MRSA XT assay |
| --- | --- |

| Prepared by | Date Adopted | Supersedes Procedure # |
| --- | --- | --- |

| Review Date | Revision Date | Signature |
| --- | --- | --- |

| Distributed to | # of Copies | Distributed to | # of Copies |
| --- | --- | --- | --- |

PRINCIPLE:

The BD MAX™ MRSA XT assay performed on the BD MAX System is an automated qualitative in vitro diagnostic test for the direct detection of methicillin-resistant Staphylococcus aureus (MRSA) DNA from nasal swabs in patients at risk for nasal colonization. The test utilizes real-time polymerase chain reaction (PCR) for the amplification of MRSA DNA and fluorogenic target-specific hybridization probes for the detection of the amplified DNA. The BD MAX MRSA XT assay is intended to aid in the prevention and control of MRSA infections in healthcare settings. It is not intended to diagnose MRSA infections nor guide or monitor treatment for MRSA infections. A negative result does not preclude nasal colonization. Concomitant cultures are necessary to recover organisms for epidemiological typing or for further susceptibility testing.

MRSA is a major cause of healthcare-acquired infections. Most transmissions occur in healthcare institutions as a result of contamination of the hands of healthcare workers, or from the healthcare environment which has been contaminated from patients carrying MRSA. While MRSA may cause infection with clinical manifestations ranging from pustules to sepsis and death1, it can also be found in the nose or on the skin of individuals (asymptomatic carriers). Treatment of MRSA infections has become a real challenge as MRSA is frequently resistant to a broad range of antimicrobial agents. Methicillin-resistant strains of S. aureus are frequently encountered in healthcare settings, and represent over 50% of hospital-acquired S. aureus isolates in some North American hospitals. Risk factors for infection with MRSA in healthcare settings include prolonged hospital stay, proximity to patients infected or colonized with MRSA, colonization with other resistant organisms such as vancomycin-resistant Enterococci (VRE) and Clostridium difficile, exposure to multiple and/or prolonged broad-spectrum antibiotic treatments, exposure to high MRSA prevalence areas within the healthcare facility, and prior MRSA infection or nasal carriage. Early identification of patients with MRSA nasal carriage can be part of an effective infection prevention program for MRSA. Culture-based detection of MRSA requires isolation of pure colonies followed by either oxacillin or cefoxitin susceptibility testing, detection of the mecA gene or detection of the penicillin binding protein (PBP 2a) encoded by the mecA gene. The culture based process takes a minimum of 24 hours with a median time to result closer to 48 hours in order to identify MRSA. With the rapidity at which MRSA infections can spread, especially in healthcare settings where carriers are common, providing MRSA nasal carriage results on the same day that the specimen was collected represents an advantage for infection prevention programs.

Active surveillance using molecular tests for rapid detection of MRSA is a proven strategy to reduce transmission in healthcare settings and prevent infection in vulnerable patients. Inaccurate detection can lead to uncontrolled transmission of MRSA and inappropriate use of healthcare resources2. With other commercial assays, up to 17.9% of positive MRSA test results are incorrect because the mecA gene is absent (commonly called “dropout mutants”)3. These false positive results can lead to unnecessary and costly isolation and treatment of patients1. Strains of MRSA with the newly discovered mecC gene account for 3-4% of all new MRSA cases4 but cannot be detected by assays that do not detect the mecC gene5. Lack of detection of the mecC gene, could lead to false negative results which may contribute to uncontrolled transmission of undetected strains of MRSA. Assay design is critical to accurate detection of  MRSA and ensure that appropriate infection control interventions are applied. The BD MAX™ MRSA XT assay uses an extended combination of primers and probes to detect new strains of MRSA, including strains with mecA or mecC gene, and decrease false positives due to mecA or mecC dropouts.

A nasal specimen is collected and transported to the laboratory using the recommended swab (refer to EQUIPMENT AND MATERIALS section). The swab is placed in a BD MAX MRSA XT Sample Buffer Tube. The Sample Buffer Tube is vortexed to release cells from the swab into the buffer. The Sample Buffer Tube is placed into the BD MAX System and the following automated procedures occur: the bacterial cells are lysed, DNA is extracted on magnetic beads and concentrated, and then an aliquot of the eluted DNA is added to PCR reagents which contain the MRSA-specific primers used to amplify the genetic targets, if present. The assay also includes a Sample Processing Control (SPC). The SPC is present in the Extraction Tube and undergoes the extraction, concentration and amplification steps to monitor for inhibitory substances as well as process inefficiency due to instrument or reagent failure. No operator intervention is necessary once the clinical sample and reagent strip are loaded into the BD MAX System. The BD MAX System automates sample lysis, DNA extraction and concentration, reagent rehydration, nucleic acid amplification and detection of the target nucleic acid sequence using real-time polymerase chain reaction (PCR). Amplified targets are detected with hydrolysis probes labeled with quenched fluorophores. The amplification, detection and interpretation of the signals are done automatically by the BD MAX System.

EQUIPMENT AND MATERIALS:

- Instrumentation / Equipment

- BD MAX™ Instrument

- BBL™ CultureSwab™ Liquid Stuart single or double swab (BD catalog no. 220099 or 220109), Copan (Venturi) Transystem™ Liquid Stuart single or double swab (Copan, catalog no. 141C or 139C)

- VWR Multi-Tube Vortexer (VWR catalog no. 58816-115)

- NALGENE® Cryogenic Vial holder (VWR catalog no. 66008-783)

- Disposable gloves, powderless

- Sterile Gauze

- Stopwatch or timer

- BD MAX™ PCR Cartridges (BD catalog no. 437519)

- Optional Materials

- Sterile scissors

- Reagents

- BD MAX™ MRSA XT Kit (BD catalog no. 443460)

Warnings and Precautions:

This test is for in vitro diagnostic use only.

- Do not use the kit if the label that seals the outer box is broken.

- Do not use reagents if the protective pouches are open or torn upon arrival.

- Close reagent protective pouches promptly with the zip seal after each use. Remove any excess air in the pouches prior to sealing.

- Do not remove desiccant from reagent pouches.

- Check reagent strips for proper liquid fills (ensure that the liquids are at the bottom of the tubes) (see Figure 1).

- Check reagent strips to ensure that all pipette tips are present (see Figure 1).

- Do not use reagents if desiccant is not present or broken inside reagent pouches.

- Do not use reagents if the foil has been opened or damaged.

- Do not mix reagents from different pouches and/or kits and/or lots.

- Do not use expired reagents and/or materials.

- Do not mix caps between tubes or re-use caps as contamination may occur and compromise test results.

- Proceed with caution when using chemical solutions as Master Mix and Extraction tube barcode readability may be altered.

- To avoid contamination of the environment with MRSA amplicons, do not break apart the BD MAX PCR Cartridge after use. The seals in the BD MAX PCR Cartridges prevent contamination.

- Performing the assay outside of the recommended time ranges may produce invalid results. Assays not performed within specified time ranges should be repeated.

- Additional controls may be tested according to guidelines or requirements of local, state, provincial and/or federal regulations or accrediting organizations.

- In cases where culture or other PCR tests are conducted in the same general area of the laboratory, care must be taken to ensure that the BD MAX MRSA XT assay, any additional reagents required for testing, and the BD MAX System are not contaminated. Gloves must be changed before manipulating reagents and cartridges.

- Always handle specimens as if they are infectious and in accordance with safe laboratory procedures such as those described in CLSI Document M296 and in Biosafety in Microbiological and Biomedical Laboratories7.

- Wear protective clothing and disposable gloves while handling kit reagents. Wash hands thoroughly after performing the test.

- Do not pipette by mouth.

- Do not smoke, drink, or eat in areas where specimens or kit reagents are being handled.

- Dispose of unused reagents and waste in accordance with country, federal, provincial, state and local regulations.

Storage and Handling Requirements:

## Specimens

- Collected specimens should be kept between 2 °C and 25 °C during transport. Protect against freezing or exposure to excessive heat.

- Specimens can be stored at 25 °C +/- 2 °C for a maximum of 48 hours or 2 – 8 °C for a maximum of 120 hours (5 days) before testing.

## Reagents

- BD MAX MRSA XT assay reagents and components are stable at 2 – 25 °C through the stated expiration date. Do not use expired components.

- BD MAX MRSA XT Master Mix and Extraction Tubes are provided in sealed pouches. To protect product from humidity, immediately re-seal after opening.

  - Reagent tubes are stable for up to 7 days at 2 – 25 °C after initial opening and re-sealing of the pouch.

  - Unreconstituted Extraction and Master Mix reagent tubes are stable for up to 5 hours at   2 – 25 ºC after being removed from their protective pouch.

PROCEDURE:

Specimen Collection and Handling

In order to obtain an adequate specimen, the procedure for specimen collection must be followed closely.

Specimen Collection/Transport/Incubation

Using a recommended swab transport device (refer to EQUIPMENT AND MATERIALS section), nasal specimens should be collected according to institutional and laboratory standard operating procedures and/or the following:

- Moisten the swab(s) with two drops (approximately 50 µL) of sterile physiological saline or use dry.

- Carefully insert the swab(s) into the patient’s nostril [a swab tip should be inserted up to 2.5 cm (1 inch) from the edge of the nares].

- Roll the swab(s) along the mucosa inside the nostril 5 times.

- Insert the same swab(s) into the second nostril and repeat steps 2 and 3.

- Replace the swab(s) in its transport tube.

- Label the transport tube.

- Transport the swab(s) to the laboratory according to hospital standard operating procedures (Refer to “STORAGE AND HANDLING REQUIREMENTS” section).

Specimen Preparation

Note: One (1) Sample Buffer Tube, one (1) Septum Cap, one (1) Master Mix (C7), one (1) Extraction Tube (B8) and one (1) Reagent Strip are required for each specimen and each External Control to be tested.

Note:  For culturing clinical specimens prior to performing the BD MAX MRSA XT assay, refer to CULTURING OF CLINICAL SPECIMENS section.

- Obtain the number of Sample Buffer Tubes corresponding to the number of specimens and external controls to be run.

- Label each Sample Buffer Tube (clear cap) with the appropriate patient identification making sure not to obscure, write, or label over the barcodes.

- Remove the cap from the Sample Buffer Tube.

- Remove the swab from the sample transport tube and place the swab in the corresponding Sample Buffer Tube.

- Hold the swab by the stem near the rim of the tube (use sterile gauze to minimize risk of contamination). Lift the swab approximately one (1) cm from the bottom of the Sample Buffer Tube and bend the stem against the edge of the tube to break it. Alternative method: use sterile scissors to cut the stem.

- Close the Sample Buffer Tube with a septum cap.

- Place Sample Buffer Tube in a NALGENE® Cryogenic Vial holder and vortex at maximum speed for one (1) minute with the Multi-Tube Vortexer. Up to 24 samples can be processed simultaneously with the Multi-Tube Vortexer.

BD MAX Operation

Note: Refer to the BD MAX System User’s Manual for detailed instructions (Operation section).

Note: The BD MAX MRSA XT assay must be performed immediately after the vortexing step above (“Specimen Preparation”, Step 7).  If retesting is necessary, re-vortex sample.

- Turn on the BD MAX System and log in by entering <user name> and <password>.

- Gloves must be changed before manipulating reagents and cartridges.

- Remove the required number of BD MAX MRSA XT Reagent Strips from the BD MAX MRSA XT Kit. Gently tap each strip onto a hard surface to ensure that all the liquids are at the bottom of the tubes.

- Remove the required number of MRSA XT Extraction Tube(s) and MRSA XT Master Mix Tube(s) from their protective pouches.  Remove excess air, and close pouches quickly with the zip seal.

- For each specimen and external control to be tested, place one (1) BD MAX MRSA XT Reagent Strip on the BD MAX System Rack, starting with Position 1 of Rack A and continuing sequentially. Do not skip spaces.

- Snap one (1) BD MAX MRSA XT Extraction Tube (white foil) into Position 1 of each BD MAX MRSA XT Reagent Strip (see Figure 1).

- Snap one (1) BD MAX MRSA XT Master Mix tube (green foil) into Position 2 of each BD MAX MRSA XT Reagent Strip (see Figure 1).

Figure 1: Snap BD MAX MRSA XT Extraction tubes and Master Mix tubes into reagent strips

- On the BD MAX software, select the <Consumable info> tab under the Run screen.

- Enter the kit lot number for the BD MAX MRSA XT assay (for lot traceability) by either scanning the barcode with the scanner or manual entry.

NOTE: Repeat steps 8 and 9 for each new lot number.

- Select the <Work List> tab, click on the <Assay> field and using the pull down menu, select <BD MAX MRSA XT>. This will automatically populate the remaining assay fields for Rack A with “BD MAX MRSA XT.”

- Enter the BD MAX MRSA XT Sample Buffer Tube ID, Patient ID and Accession Number (if applicable) for Position 1 of Rack A, either by scanning the 1D barcode with the scanner or by manual entry.

- Click on the <Lot Number> field and using the pull down menu, select the kit lot number (on the outer box). This will automatically populate the remaining lot number fields for Rack A with the same lot number.

- Enter information for Position 2 of Rack A and continue for all remaining Sample Buffer Tubes in the rack.

NOTE: Steps 12 and 13 must be repeated for each new lot number.

- Repeat steps 10 to 13 for Rack B.

- Place the BD MAX MRSA XT Sample Buffer Tube(s) in the BD MAX Rack(s) following the same order as entered in the worklist.  Do not skip or leave empty positions between tubes.

NOTE: Place tubes into the sample rack with 1D barcode labels facing outward (this makes scanning tubes easier during sample login).

- Place the required number of BD MAX PCR Cartridge(s) into the BD MAX System (see Figure 2).

- Each cartridge Accommodates 2 runs of up to 12 samples for a total of 24 samples.

- The BD MAX System will automatically select the position and row on the PCR cartridge for each run.

- Cartridges are used on a per-run AND rack basis (2 runs per cartridge and 1 cartridge per rack).

Figure 2:  Load PCR Cartridges

- Load Rack(s) into the BD MAX System (Figure 3). Ensure that the placement of rack(s) (left to right) corresponds to the worklist created (top to bottom).

Figure 3:  Load Rack(s) into the BD MAX™ System.

- Close the BD MAX System lid and click the <Start Run> button to begin processing.

- At the end of the run, check results immediately or store Sample Buffer Tubes at 2-8 °C until the results are checked.

NOTE: If a septum was damaged during the run, replace it with a new one before storing the specimen.

NOTE: Sample Buffer Tubes can be stored at 25 °C +/- 2 °C for a maximum of 36 hours or at 2-8 °C for a maximum of 120 hours (5 days) after the run has been started. When an Indeterminate (IND), Unresolved (UNR), or Incomplete (INC) result is obtained, or when an External Control failure occurs, a repeat test from the Sample Buffer Tube must be performed within this timeframe (see REPEAT TEST PROCEDURE section).

# QUALITY CONTROL:

Quality control procedures monitor the performance of the assay. Laboratories must establish the number, type and frequency of testing control materials according to guidelines or requirements of local, provincial, state and/country regulations or accreditation organizations. For general QC guidance, the user may wish to refer to CLSI MM38 and EP129.

1.	An External Positive Control is intended to monitor for substantial reagent failure while an External Negative Control is used to detect reagent or environmental contamination (or carry-over) from other specimens or MRSA amplicons. External Control materials are not provided by BD. Various types of External Controls are recommended to allow the user to select the most appropriate for their laboratory quality control program:

- Commercially available control materials [e.g., a reference MRSA strain (ATCC™* 43300) can be used as positive control. Staphylococcus epidermidis strain (e.g., ATCC 12228) can be used as negative control].

- Previously characterized specimens known to be positive or negative for MRSA.

NOTE: It is recommended that bacterial strains be freshly prepared in saline to a turbidity of 0.5 McFarland (~1.0 X 108 CFU/mL) from isolated colonies and subsequently diluted with saline to obtain a final concentration of ~1.0 X 104 CFU/mL. Dip a swab into the diluted bacterial suspension, express the excess liquid, place the swab in a corresponding Sample Buffer Tube, and follow instructions described at step 5 section of “SPECIMEN COLLECTION / TRANSPORT / INCUBATION” section.

2.	One (1) External Positive Control and one (1) External Negative Control should be run daily until adequate process validation is achieved on the BD MAX System. Reduced frequency of control testing should be based on adequate data and determined by the individual laboratory.

3.	An External Negative Control that yields a positive test result is indicative of a specimen handling and/or contamination problem. Review the specimen handling technique to avoid mix-up and/or contamination. An External Positive Control that yields a negative result is indicative of a specimen handling/preparation problem. Review the specimen handling/preparation technique.

4.	An External Control that yields an Unresolved, Indeterminate or Incomplete test result is indicative of a reagent or a BD MAX System failure. Check the BD MAX System monitor for any error messages. Refer to the "System Error Summary" section of the BD MAX System User’s Manual10 for interpretation of warning and error codes. If the problem persists, use reagents from an unopened pouch or use a new BD MAX MRSA XT assay kit.

NOTE: External Positive and Negative Controls are not used by the BD MAX System software for the purpose of sample test result interpretation. External Controls are treated as if they were patient samples.

- Each BD MAX MRSA XT assay Extraction Tube contains a Sample Processing Control (SPC) which is a plasmid containing a synthetic target DNA sequence. The SPC will be extracted, eluted and amplified along with any DNA present in the processed specimen, ensuring the predictivity of the assay. The SPC monitors the efficiency of DNA capture, washing and elution during the sample processing steps, as well as the efficiency of DNA amplification and detection during PCR analysis. If the SPC result fails to meet the acceptance criteria, the result of the specimen will be reported as Unresolved. An Unresolved result is indicative of specimen-associated inhibition or processing or reagent failure. Repeat any specimen reported as Unresolved according to the "REPEAT TEST PROCEDURE" section below.

RESULTS INTERPRETATION

Results are available on the Results tab in the Results window on the BD MAX System monitor. The BD MAX System software automatically interprets test results. A test result may be called as MRSA NEG (negative), MRSA POS (MRSA positive) or MRSA UNR (Unresolved) based on the amplification status of the target and of the Sample Processing Control. IND (indeterminate) or INC (incomplete) results are due to BD MAX System failure. Results are based on the following decision algorithm.

Table 1: BD MAX MRSA XT assay Decision Algorithm

| ASSAY RESULT REPORTED | INTERPRETATION OF RESULT |
| --- | --- |
| MRSA POS | MRSA DNA detected |
| MRSA NEG | No MRSA DNA detected |
| MRSA UNR | Unresolved result. No target amplification; no SPC amplification |
| IND | Indeterminate result due to BD MAX System failure (with Warning or Error Codes*) |
| INC | Incomplete Run (with Warning or Error Codes*) |

* Refer to the “Troubleshooting” section of the BD MAX System User’s Manual for interpretation of warning and error codes.

MRSA POS (MRSA DNA detected)

- Fluorescence signal is detected for both MREJ (S. aureus specific) and mecA or mecC targets.

- The SPC is ignored since MRSA target amplification overrides this control.

MRSA NEG (no MRSA DNA detected)

- Fluorescence signal is not detected by the BD MAX MRSA XT assay for any target (mecA, mecC or MREJ targets) and fluorescence signal is detected for the SPC; or

- Fluorescence signal is detected for the mecA or mecC gene only [The mecA and mecC genes are not unique to S. aureus species and can be found in other bacterial genera (e.g. S. epidermidis)]; or

- Fluorescence signal is detected for MREJ only (mecA or mecC gene drop-out).

MRSA UNR (Unresolved result)

- Fluorescence signal is not detected for mecA, mecC or MREJ targets; and

- Fluorescence signal not detected for the SPC (inhibitory specimen or reagent failure).

IND (Indeterminate result)

- BD MAX System failure with Warning or Error Codes. Refer to the “Troubleshooting” section of the BD MAX System User’s Manual10 for interpretation of warning and error codes.

INC (Incomplete run)

- BD MAX System failure with Warning or Error Codes. Refer to the “Troubleshooting” section of the BD MAX System User’s Manual10 for interpretation of warning and error codes.

REPEAT TEST PROCEDURE

NOTE: Sufficient volume is available for only one repeat test from the Sample Buffer Tube on the BD MAX System. For Sample Buffer Tubes stored at 25 °C +/- 2 °C, retesting must be performed within 36 hours of the steps covered in the “SPECIMEN PREPARATION” section above. Alternatively, for Sample Buffer Tubes stored at 2 – 8 °C, retesting must be performed within 120 hours (5 days) of the steps covered in the “SPECIMEN PREPARATION” section above.

NOTE: New samples may be tested in the same run with repeat samples.

UNRESOLVED RESULT

Unresolved results may be obtained in the event that an inhibitory substance prevents proper target, or SPC amplification. Sample(s) can be repeated from their corresponding Sample Buffer Tube(s) within the timeframe defined above. Vortex the sample(s) for one (1) minute and restart from the “BD MAX Operation” section.

INDETERMINATE RESULT

Indeterminate results may be obtained in the event that a System failure occurs. Sample(s) can be repeated from their corresponding Sample Buffer Tube(s) within the timeframe defined above. Vortex the sample(s) for one (1) minute and restart from the “BD MAX Operation” section. For the interpretation of warning or error code messages, refer to the BD MAX Software User’s Manual10 (“Troubleshooting” section).

INCOMPLETE RESULT

Incomplete results may be obtained in the event that the Sample Preparation or the PCR did not reach its expected time points. Sample(s) can be repeated from their corresponding Sample Buffer Tube(s) within the timeframe defined above. Vortex the sample(s) for one (1) minute and restart from “BD MAX Operation” section. For the interpretation of warning or error code messages, refer to the BD MAX System User’s Manual10 (“Troubleshooting” section).

EXTERNAL CONTROL FAILURE

External Controls should yield expected results when tested. If specimens have to be repeated due to an incorrect External Control result, they should be repeated from their Sample Buffer Tube along with freshly prepared External Controls within the timeframe defined above. Vortex the samples for one (1) minute and restart from the “BD MAX Operation” section.

CULTURING OF CLINICAL SPECIMENS

In order to perform antimicrobial susceptibility testing or epidemiological typing, clinical specimens may be cultured from the collection device (swab) prior to performing the specimen preparation procedure (using a Streak-Plate method) or after the specimen preparation procedure (using an Enrichment Broth method). Immediately after the end of the initial PCR run, swabs may be stored at 2-8 °C for up to 36 hours in Sample Buffer Tubes before culturing, following hospital procedures.

LIMITATIONS OF THE PROCEDURE

- This product is intended for use with nasal swab specimens collected using specimen collection and transport devices listed in the EQUIPMENT AND MATERIALS section. Performance of the BD MAX MRSA XT assay using Liquid Amies single or double swab transport device has not been established.

- This product should only be used with the BD MAX System.

- Incorrect test results may occur from improper specimen collection, handling or storage, technical error, sample mix-up or because the number of organisms in the specimen is below the analytical sensitivity of the test. Careful compliance with the package insert instructions and the BD MAX System User’s Manual10 are necessary to avoid erroneous results.

- Good laboratory technique is essential for the proper performance of this assay. Due to the high analytical sensitivity of this test, extreme care should be taken to preserve the purity of all materials and reagents.

- Screening determines the colonization status at a given time. Colonization may vary depending upon patient treatment (e.g. decolonization regime), patient status (e.g. transient MRSA colonization) or exposure to high-risk environments (e.g. contact with MRSA carrier and/or prolonged hospitalization). Colonization status should be monitored according to institutional policies.

- A BD MAX MRSA XT positive result does not necessarily indicate eradication treatment failure since DNA presence may persist. A negative result following a previously positive test result may indicate eradication treatment success or may occur due to intermittent colonization.

- A positive test result does not necessarily indicate the presence of viable organisms. A positive result is indicative of the presence of MRSA DNA. The BD MAX MRSA XT assay simultaneously detects the mecA or mecC gene carried within the SCCmec cassette and a S. aureus specific sequence located within the junction of the SCCmec cassette and the orfX gene (MREJ).

- The BD MAX MRSA XT assay is designed to detect MREJ genotypes i, ii, iii, iv, v, vi, vii, ix, xiii, xiv and xxi which represents most of mecA and mecC harboring MRSA strains (belonging to different SCCmec/MREJ types) accounting for more than 98% of worldwide strains tested by BD to date. The ability of BD MAX MRSA XT assay to detect other MREJ genotypes is unknown.

- The BD MAX MRSA XT assay does not report Borderline Oxacillin Resistant S. aureus (BORSA) as MRSA (it will report as NEG). The mechanism of oxacillin resistance in BORSA strains is due to an increased production of ß-lactamases, not the mecA or mecC gene. BORSA strains are rare.

- The BD MAX MRSA XT assay performance in detecting modified S. aureus (MOD-SA) is not known as those strains have not been evaluated. The mechanism of oxacillin resistance in MOD-SA strains is due to changes in affinity of penicillin-binding proteins for oxacillin. MOD-SA strains are rare.

- The BD MAX MRSA XT assay will generate a false positive MRSA result when testing a co-colonized nasal specimen containing both a methicillin-resistant coagulase negative Staphylococcus (MRCoNS) and an "empty cassette" methicillin-susceptible SA variant. Co-colonization with MRCoNS and an “empty cassette” methicillin-susceptible SA is rare.

- As with all PCR-based in vitro diagnostic tests, extremely low levels of target below the LoD of the assay may be detected, but results may not be reproducible.

- Tobramycin may interfere with the BD MAX MRSA XT assay (refer to “Interfering Substances” section for further details).

- False negative results may occur due to loss of nucleic acid from inadequate collection, transport or storage of specimens, or due to inadequate bacterial cell lysis. The Sample Processing Control has been added to the test to aid in the identification of specimens that contain inhibitors to PCR amplification and as a control for reagent integrity and of the assay system as a whole. The Sample Processing Control does not indicate if nucleic acid has been lost due to inadequate collection, transport or storage of specimens, or if bacterial cells have been adequately lysed.

- In a mixed culture, the detection of MRSA is variable when high concentrations of MRSE are present. Competition from MRSE was observed at an MRSA:MRSE ratio of 1: ≥ 1x103.

- In a mixed culture, the detection of MRSA is variable when high concentrations of MSSA are present. Competition from MSSA was observed at an MRSA:MSSA ratio of  1: ≥  1x104.

- BD MAX MRSA XT assay results may sometimes be Unresolved due to an invalid Sample Processing Control, or be Indeterminate or Incomplete due to instrument failure, and require retesting that can lead to a delay obtaining final results.

- Mutations or polymorphisms in primer- or probe-binding regions may affect detection of new or unknown MRSA, resulting in a false negative result with the BD MAX MRSA XT assay.

- As with all in vitro diagnostic tests, positive and negative predictive values are highly dependent on prevalence. BD MAX MRSA XT assay performance may vary depending on the prevalence and population tested.

- The BD MAX MRSA XT assay requires use of three (3) optical channels from the BD MAX System; 475/520 channel, 585/630 channel and 680/715 channel. Performance of the remaining optical channel has not been established with this assay.

# INTERFERING SUBSTANCES

Twenty nine (29) microorganisms and chemical substances which might be used in the nares or found in nasal swab specimens were evaluated for potential interference with the BD MAX MRSA XT assay (Table below). MRSA negative and MRSA positive samples at 2-3 x LoD were tested with the highest amount of each compound or microorganism likely to be found at the sampling site or on the nasal swab sample. Results demonstrated no reportable interference with any microorganisms or chemical substance except for Tobramycin which showed interference with the BD MAX MRSA XT assay when tested at a concentration of 4.5 x 10-3 g/swab.

Endogenous and Commercial Exogenous Substances Tested with the BD MAX MRSA XT assay

| Substance | Result |  | Substance | Result |
| --- | --- | --- | --- | --- |
| Mucin, from bovine submaxillary glands | NI |  | Rhinocort aqua®* | NI |
| Dexamethasone Sodium Phosphate Ophtalmic Solution USP, 0.1% Dexamethasone Phosphate Equivalent | NI |  | Zicam®* No-Drip Liquid Nasal GelTM* Extreme Congestion Relief | NI |
| Chloraseptic®* | NI |  | Fluticasone Propionate | NI |
| Taro-Mupirocin, Mupirocin Ointment USP, 2% | NI |  | Luffeel®* | NI |
| Long Lasting Dristan®* Nasal Mist | NI |  | Staphylococcus epidermidis | NI |
| Neo-Synephrine®* | NI |  | Micrococcus luteus | NI |
| Equate® Nasal Spray Decongestant | NI |  | Enterococcus faceium | NI |
| Beconase AQ®* | NI |  | Enterococcus faecalis | NI |
| Flunisolide Nasal Solution USP, 0.025% | NI |  | Escherichia coli | NI |
| Nasacort®* AQ | NI |  | Corynebacterium flavescens | NI |
| Nasonex®* | NI |  | Moraxella catarrhalis | NI |
| Relenza®* | NI |  | Staphylococcus hominis subsp. hominis | NI |
| Tobramycin | I |  | Haemophilus influenza | NI |
| Blood | NI |  | Streptococcus pneumoniae | NI |
| FluMist®* | NI |  |  | NI |

NI: No reportable interference with the BD MAX MRSA XT assay.

I: Reportable interference with the BD MAX MRSA XT assay.

PRECISION:

Within-laboratory precision was evaluated for the BD MAX MRSA XT assay at one (1) site. The Precision panel consisted of 4 sample categories near the LoD. Each specimen contained simulated nasal matrix. MRSA strains were tested as follows:

- Moderate Positive (MP) (MRSA MREJ Type ii): ≥ 2 and ≤ 5 x LoD

- Low Positive (LP) (MRSA MREJ Type ii): ≥ 1 and < 2 x LoD

- Low Positive (LP) (MRSA MREJ Type vii): ≥ 1 and < 2 x LoD

- High Negative (HN) (MRSA MREJ Type ii): < 1 x LoD

- True negative (TN): Negative specimen (no target)

Testing was performed in duplicate, over 12 days, with 2 runs per day, by 2 different technologists. Precision study results for TN, MP, LP, and HN MRSA samples demonstrated 100%, 100%, 97.9%, and 60.4% agreement, respectively.

## REFERENCES:

- Blanc et al., High Proportion of Wrongly Identified Methicillin-Resistant Staphylococcus aureus Carriers by Use of a Rapid Commercial PCR Assay Due to Presence of Staphylococcal Cassette Chromosome Element Lacking the mecA Gene. J. Clin. Microbiol 2011;49:722-724.

- Jarvis et al., National prevalence of methicillin-resistant Staphylococcus aureus in inpatients at United States health care facilities, 2010. Am. J Infect Control 2012; 40:194-200.

- Farlay et al., Comparison of the BD GeneOhm Methicillin-Resistant Staphylococcus aureus (MRSA) PCR Assay to Culture by Use of BBL CHROMagar MRSA for Detection of MRSA in Nasal Surveillance Cultures from an At-Risk Community Population, J. Clin. Microbiol.2008;46:743-746.

- Petersen et al., Epidemiology of methicillin-resistant Staphylococcus aureus carrying the novel mecC gene in Denmark corroborates a zoonotic reservoir with transmission to humans. Clin Microbiol Infect. 2013;19:E16-E2

- Shore et al., Detection of Staphylococcal Cassette Chromosome mec Type XI Carrying Highly Divergent mecA, mecI, mecR1, blaZ, and ccr Genes in Human Clinical Isolates of Clonal Complex 130 Methicillin-Resistant Staphylococcus aureus, Antimicrob. Agents and Chemother.  2011;55:3765-3773

- Clinical and Laboratory Standards Institute. Protection of laboratory workers from occupationally acquired infections; Approved Guideline. Document M29 (Refer to the latest edition).

- Centers for Disease Control and Prevention, and National Institutes of Health. Biosafety in microbiological and biomedical laboratories. Chosewood L.C. and Wislon D.E. (eds) (2009). HHS Publication No. (CDC) 21-1112.

- Clinical and Laboratory Standards Institute. Molecular Diagnostic Methods for Infectious Diseases; Approved Guideline, Document MM3 (Refer to the latest edition).

- Clinical and Laboratory Standards Institute. User Protocol for Evaluation of Qualitative Test Performance; Approved Guideline, Document EP12 (Refer to the latest edition).

- BD MAX System User’s Manual (refer to the latest version) BD Diagnostics, Sparks, MD, USA.

_____________________________________________________________________________________________

The purchase of this product allows the purchaser to use it for amplification and detection of nucleic acid

sequences for providing human in vitro diagnostics. No general patent or other license of any kind other

than this specific right of use from purchase is granted hereby.

This product is sold under license, and purchase of this product does not include rights to use for certain

blood and tissue screening applications, nor for certain industrial applications.

BD Diagnostics Technical Service 1.800.638.8663

Geneohm Sciences Canada, Inc., 2555 Boul. du Parc Technologique, Québec (QC), G1P aihf    4S5, Canada

Benex Limited, Pottery Road, Dun Laoghaire, Co. Dublin, Ireland

*Brands are trademarks of their respective owners.

BD and BD logo and all other trademarks are property of Becton, Dickinson and Company. © 2014 BD

Made in Canada.

# Footnotes

This “Sample Procedure” is not intended as a substitute for your facility procedure manual, instrument manual, or reagent labeling/package insert. This “Sample Procedure” is intended as a model for use by your facility to be customized to meet the needs of your laboratory.