# TTS WIRING CONNECTIONS

# Redwood Sciences Lab

USDA Forest Service

# Rand Eads, Jack Lewis

# Revised  10/30/2006

# Use with TTS Program Revisions 4.0 – 5.0

### Configuration:

### Campbell CR510 Data Logger

### DTS-12 Turbidity Sensor

## Campbell CR510	Wire				Description

## POWER IN (12V)	Red				Bat pos. (see Note 1)

POWER IN (G)	Blk				Bat neg.

EARTH GROUND	Grn				Ground Rod

#### ISCO Power Connection

12V			Red				See Note 2

G			Blk				ISCO Ground

Druck 1230-1830 Pressure Transducer with DES2 (use the wiring

table from the Druck manual if not using the desiccant box)

H1 DIFF		Yel				Output +

L1 DIFF		Blu				Output -

AG			Blk				Supply -

E1			Red				Supply +

G			Clr				Shield

DTS-12 Turbidity Probe

C2			Wht				DTS-12 Data

12V			Red				DTS-12 Power

G			Blk				DTS-12 Ground

ISCO Flow Connection (see Note 5)

C1			Red				ISCO C

G			Blk				ISCO B

Campbell T107 Air Temperature (optional)

SE4			Red				Signal

E2			Blk				Supply +

AG			Pur				Supply –

G			Clr				Shield

Tipping Bucket Rain Gage (optional)

P1			Red				------

G			Wht				------

Note 1: Use an in-line automotive fuse holder and 10 amp fuse between the 12V gel cell battery and CR10X.

Note 2: The ISCO power +12V, battery +12V, and Solar +12V (if connected) must be soldered together before the diode in the 9V Battery Backup diagram.

Note 3: The instruction in the TTS program will only work with the Campbell T107 temperature probe (not the T108).

##### Note 4: Water temperature measurements can be collected with the DTS-12 sensor.

Note 5: Former incarnations of this document (prior to Oct 2006) documented two options for the ISCO flow connection. One required a Crydom relay and the other required CSI control cable 10164-L.  We now recommend the simpler option shown in the table above. The user must purchase a flow cable from ISCO (for a non-ISCO flow meter) or build a cable from parts.

Note 6: Always confirm that the wire color/description in this document agrees with manufacturer diagram or table.  Wire color-coding is manufacturer-specific and, in some cases, may not conform to the table above