Institute of Space

Technology Islamabad

Procurement Department

LIT/IST Div/1718/3000000104

07-Jun-2018

To:    _____________

__________________

Subject: Invitation of Tender for Fabrication of Ambulance (Toyota Hiace)

Dear Sir,

We invite you to quote for supply of the items as detailed in Annex-A & B enclosed. Offers must be submitted on company’s letterhead with complete specifications of offered equipment/items and undertaking given at the end of Annex-A, B & D along with format for bid (Annex-C) to be duly filled and signed.

The terms and conditions of this tender are as under:-

- The bids comprise a single package containing two separate sealed envelopes are required to be delivered at IST-Islamabad latest by 11:30 hours on 
22-Jun-2018. Tenders will be opened on same date at 12:00 hours in the presence of bidders who desire to attend. Each envelope shall contain separately the financial proposal and the technical proposal. The envelopes shall be marked as “FINANCIAL PROPOSAL” and “TECHNICAL PROPOSAL” in bold and legible letters to avoid confusion. Initially, only the envelope marked “TECHNICAL PROPOSAL” shall be opened. The envelope marked as “FINANCIAL PROPOSAL” shall be retained in the custody of the procuring agency without being opened. The financial proposals of bids shall be opened publicly at a time, date and venue announced and communicated to the bidders in advance.

- Sealed tenders addressed to the Assistant Director (Procurement), Institute of Space Technology, Near Rawat Toll Plaza, Islamabad Highway, Islamabad may be dropped in the tender box or mailed through courier, registered post; but it must be ensured that the same reaches our office well in time as late offers will not be considered. The envelopes should be marked clearly with the following information.

Tender No	:	LIT/IST Div/1718/3000000104

Due Date		:	22-Jun-2018

Offer			:	Technical/Financial

Tender for	:	Procurement of Fabrication of

Ambulance

- Rates:	The bidder should quote their rates for:-

- DDP/ Free Delivery IST inclusive of all Govt. taxes. Details of such taxes to be provided.

- The bidder should indicate their NTN NO. & Sales Tax registration number.

- Items must be quoted as per user requirements/sample available at IST - Store. Before submitting any offer, bidder must consult the concerned user regarding any query/specifications or checking of sample. No change in the offer will be accepted later on.

- Offer shall be made as per our demanded specs or equivalent.

- Bid Security: Quotations should be accompanied with 02% of total bid value in the form of pay order/ call deposit/ Demand draft in favor of IST-Islamabad. No offer will be entertained without bid security. The bid security will finally be adjusted/ refunded on completion of the contract. At any stage if the purchaser feels that the bidder has provided wrong information, fake document or the sample is not provided as per the specifications mentioned in their offer, the earnest money will be forfeited. (Bid Security must be enclosed in separate sealed envelope clearly marked “Bid Security”).

- IST FTN #: 9010613-0.

- Delivery: Items are required to be delivered within 30 days or earlier after issuance of Purchase Order. Delivery of the ordered items at the Purchaser’s premises will be Supplier’s responsibility.

- Validity: The offer shall remain valid for our acceptance for 90 Days from the date of opening of tender.

- Pre-Shipment Inspection:  Pre-Shipment inspection for 2 x IST officials covering all expenses i.e travel, accommodation, DA, 15% incidental charges of DA as per govt. TA/DA rules.

- Inspection: The inspection shall be carried out at IST by IST technical committee and the inspection report (findings) will be binding on you.

- Taxes:

- GST/CED will be paid by the PURCHASER to the SUPPLIER’S on production of a separately numbered ‘Tax Invoice” showing the amount of sales tax and other details under section 23 of Sales Tax Act-1990.

- As per in vogue GST SRO, 20% of the 17% GST amount will be retained by the Purchaser and later on submitted in the CBR / Government Treasury, where as 80% of the 17% GST will be paid to the Supplier for onward deposition in CBR/Government Treasury.

- Income tax will be deducted as per rule.

- Any change in taxes/ duties levied by Govt. after opening of bids will be on account of bidder. No change in offer will be accepted later on.

- Payment:  All payments will be made through crossed cheque in favor of the SUPPLIER in Pak Rupees. The payment mode will be “100 % Payment after Delivery” after satisfactory performance in every aspect and inspection.

Note: Any representation for the delay of payment will not be accepted, until complete required documents have not been received. (Countdown starts when Store receives complete documents). Advance Payment can only be made against submission of equivalent amount of Bank Guarantee.

- Late Offers: The quotations received late (after opening the tender) will not be considered. Cutting/over writing is not allowed, unless/until authenticated under proper signature of the authorized person.

- Performance Bond: IST may ask to provide Performance Bank Guarantee equivalent to 10% of total order value. The Performance Bond shall be valid till the Guarantee/Warranty period of particular supply.

- IST reserves the right to increase/decrease the quantity as per requirements.

- Publication/Literature/Catalogue in Hard Copies/CDs will be provided along with offer. Moreover the supplier will be bound to provide the following.

- Essential Running Spares to be provided.

- Certificate of Test Data/Calibration will be provided.

- Supplier should be authorized partner or reseller of OEM.

- Quality Standards Certifications must be provided that include:

- Quality Certification for Firm as per ISO (ISO-9001, ISO-14001, OHSAS 18001) required.

- CE Standards (preferably EN 1782/KKK-A-1822F) Certification of quoted products / medical equipment required.

- Supplier shall ensure to supply the stores of high quality and give warranty of its satisfactory function, good workmen-ship and as per required specifications. Supplier will be required to provide at least Standard Warranty/ Guarantee for the supplied stores.

- Installation and Commissioning: Supplier offer must be inclusive of complete installation and commissioning of supplied items / equipment from Firm’s Factory premises. The supplied equipment must be made 100% operative up to the satisfaction of user.

- Bids will be evaluated in a manner that the items will bring “value for money” to the organization. Offers with lower specifications than demanded specs will be liable for rejection.

- If IST deemed necessary, the lowest bidder may be asked to provide/preview finished sample before awarding of any contract/order. In case the sample of lowest bidder is rejected, sample from other bidders will be considered and contract will be awarded to technically qualified lowest bidder.

- Conditional bids may be liable for rejection. Offers with incomplete documentation will be rejected.

- IST may ask the lowest bidder to provide a “Certificate” that firm has never been black listed which will be submitted on stamp paper of valuing Rs. 20/-.

- Failure & Termination: Incase the supplier / contractor failed to deliver the   stores within specified delivery period, IST shall have the following options:-

- To recover liquidated damages @ 2% of total purchase order value per month or part thereof will be imposed. The maximum limit of liquidated damages will not exceed 10% of contract value. Such damages shall be recovered from the bills, earnest money and or by withholding any amount due to be paid against this or any other purchase order.

- To purchase from elsewhere, without notice to you but at your risk and cost, the stores not delivered within the stipulated period and to recover excess value so paid by us from your bills or from your security and or to cancel the contract at your risk and cost.

- Both technical and financial offers must be submitted in two separate sealed envelopes. Each envelope should contain the documents as under:

| Technical offer: Technical quotation on letter head without prices. Format of bid as per Annexure-C (attached). Complete technical catalogues/CDs/broachers etc. | Financial offer: i. Quotation on letter head with prices. ii. 02 % bid security (copy must be attached with technical bid and its reference no. filled in Annex C) |
| --- | --- |

- Blacklisting:	If the bidder/supplier offered wrong/fake information in bid or  failed to execute order as per laid down terms and conditions in the tender/order or the performance is not satisfactory or are found to be indulging in corrupt or fraudulent practices, the action for blacklisting of firm/ supplier/ contractor will be taken as per PPRA rule 19. In case of any dispute, the case will be referred to Vice Chancellor (IST), whose decision will be final.

- IST reserves the rights to accept or reject any or all tenders as a whole or in part as per PPRA Rule 33 (1). The decision in this regard will be firm, final and binding on all bidders.

(Salman Qasim)

Asstt Director (Procurement)

Institute of Space Technology, Islamabad

APPENDIX-A

| The Following Requirements are mandatory. Sr. # | Qualification Requirements / Knock Down Criteria | Checklist |
| --- | --- | --- |
| 1 | Certificate of Company/Firm Registration/Incorporation under the laws of Pakistan | M |
| 2 | Copy of CNIC | M |
| 3 | Proof of Sales Tax Registration, NTN Certificate (Status = Active with FBR) | M |
| 4 | Relevant Business should be active for at least 2 years (Should attach POs) | M |
| 5 | Undertaking that the firm is not blacklisted and involved in litigation with Government (Affidavit duly attested by the Oath Commissioner/Notary Public) | M |
| 6 | List of satisfied customers along with their contact numbers | M |
| 7 | Bidder shall submit an affidavit worth Rs. 50/-/ or Company Letter Head that he accepted all terms and conditions of bidding documents and PPRA rules 2004 and shall continue the same in future, duly signed and stamped | M |
| M = Mandatory |
| All mandatory documents MUST be submitted with the Bid. Failure to fulfill any of the above condition will disqualify the bidder from the bidding process. Submission of all documents mentioned above is mandatory at the time of the tender submission. No document will be accepted on later stage. |

Annexure – A

Technical Offer

LIT/IST Div/1718/3000000104

Date: 07-Jun-2018

Due On: 22-Jun-2018

| Installation Work (Fabrication of Ambulance Vehicle): |
| --- |
| S/No | Description | Qty | Yes | No | Alternative Offer |
| 1 | FLOOR FLOOR WOULD NOT BE PUNCTURED FOR INSTALLATION. FLOOR MADE OF WATER AND TERMITE PROOF HALF INCH THICK PLY-WOOD. THE FLOOR WOULD BE CARPETED WITH VINYL SHEET 2-3 MM THICK. THE FLOORING WOULD COVER ALL SIDES AND GAPS OF THE CABINETRY MAKING IT COMPLETELY WATER TIGHT. SPECIAL HARD POINTS OF 5MM THICKNESS MS SHEET MUST BE CREATED FOR INSTALLATION OF CHAIR, CABINET AND BENCH ETC. A SPECIAL S/S 16SWG SHEET WILL BE PROVIDED ON THE FLOOR UNDER THE STRETCHER AREA. | 1 |  |  |  |
| 2 | CEILING MADE OF FG MATERIAL WITH INTERNAL LED LIGHTS, ONE LONG GRIP HANDLE ALONG THE LENGTH OF CEILING MADE OF POWDER COATING PROPERLY REINFORCED AND OPTION FOR HOLDING TWO IV HOOKS WITH ANTI-SWING BELTS IN CEILING OR SIDE PANEL | 1 |  |  |  |
| 3 | PARTITION WALL PARTITION WALL BETWEEN THE DRIVER AND PATIENT COMPARTMENT WOULD BE MADE OF FG/ABS SHEET FIXED WITH STEEL BRACKETS TO SIDE WALLS. A SMALL WINDOW WOULD BE GIVEN IN BETWEEN THE DRIVER AND PATIENT COMPARTMENT WITH A BUILT-IN PLACE WITH GLASS FOR COMMUNICATION | 1 |  |  |  |
| 4 | SIDE UTILITY VERTICAL PANEL MADE OF ABS/FG MOLDED MATERIAL ON THE RIGHT SIDE DEDICATED O2 SYSTEM HAVING TWO OUTLETS, O2 ALARM SYSTEM, O2 CYLINDER CHANGEOVER SYSTEM AND ONE PRESSURE GAUGE FOR LINE PRESSURE. ALL SWITCHES AND SOCKETS FOR 12 V & 220 V, HAVING 220 INVERTER AT THE BACK | 1 |  |  |  |
| 5 | SIDE CABINET FOR EQUIPMENT MADE OF COMPLETELY FG MATERIAL ON RIGHT SIDE OF THE VEHICLE STARTING FORM THE END OF MAIN CABINET TO THE FULL LENGTH OF RIGHT SIDE, ONE TOP SHELF TO ACCOMMODATE AED, MONITOR, EMERGENCY BAG, SEARCH LIGHTS, ALL HAVING SPECIAL BRACKETS; LOWER SHELF HAVING SUCTION UNIT, ONE WIRED SHELF FOR ACCESSORIES AND OTHER ITEMS | 1 |  |  |  |
| 6 | ELECTRICAL WIRING FIRE RETARDANT AND HIGH TEMPERATURE RESISTANT ALL ELECTRICAL WIRING OF AMBULATORY COMPARTMENT WOULD BE OF WIRING HARNESS SYSTEM WITH NO JOINT AND TAPING HAVING ALL CONNECTOR SYSTEM AND HAVING A SEPARATE FUSE BOX FOR EACH OPERATION. ALL WIRES WOULD BE COLOR COATED AND WOULD BE OF FIRE RETARDANT AND HIGH TEMPERATURE RESISTANT MATERIAL. THERE WOULD BE A SEPARATE MAINTENANCE-FREE BATTERY OF 40 AMPS FOR AMBULATORY COMPARTMENT WHICH IS TO BE CHARGED THROUGH THE VEHICLE GENERATOR. THERE WOULD BE A 220 VOLTS INLET AND OUTLET SOCKETS FOR OUTSIDE SOURCE VOLTS | 1 |  |  |  |
| 7 | THREE TO FOUR PERSON BENCH MADE OF 3” SPECIAL POLYURETHANE FOAM AS USED IN AUTOMOBILE SEATS WITH APPROXIMATE 16” WIDTH, MADE OVER A STORAGE BOX OPEN-ABLE IN TWO PARTS. HEAT SEALED SEAT COVER OF BEST QUALITY ARTIFICIAL LEATHER, SEAT BELTS AND CUSHIONED BACK. ONE SIDE BAR MADE OF 1” M.S POWDER COATED PIPE | 1 |  |  |  |
| 8 | ONE PERSON CREW SEAT ONE CREW SEAT FIXED TO THE BODY OF THE VEHICLE WITH WAIST SEAT BELT AT THE HEAD END OF PATIENT WITH BEST QUALITY ARTIFICIAL LEATHER / FABRIC | 1 |  |  |  |
| 9 | OXYGEN SECTION TO ACCOMMODATE TWO M.F SIZE O2 CYLINDERS, WOULD HAVE DOUBLE STEEL BRACKETS WITH RUBBER PADDING. AN INTERNAL LIGHT WITH A DOOR SWITCH. FLOOR OF CABINET MATTED WITH VINYL FLOORING | 1 |  |  |  |
| 10 | WASH BASIN AND DUST BIN INTEGRATED WASH BASIN WITH WATER TAP AND WATER RESERVOIR WITH ELECTRIC WATER PUMP AND ELECTRIC SWITCH FOR TAP. DUST BIN IN THE CABINET | 1 |  |  |  |
| 11 | DRAWER SECTION HAVING 3 TO 4 DRAWERS. ALL DRAWERS MADE OF S/S SHEET MATERIAL. ALL DRAWERS HAVING SPECIAL SELF CLOSING & SELF RETAINING CATCHER HANDLE LOCK SYSTEM WHICH IS NOT SELF OPEN-ABLE DURING EMERGENCY DRIVING | 1 |  |  |  |
| 12 | SPOT LAMP SPOT LAMP WITH FLEXIBLE NECK FOR PATIENT EXAMINATION | 1 |  |  |  |
| 13 | BACK FLOOD LIGHT BACK SEARCH LIGHT, ADEQUATE INTERNAL LIGHTING FOR HANDLING THE PATIENTS AT NIGHT | 1 |  |  |  |
| 14 | EXTERNAL AND INTERNAL LIGHT AND WARNING DEVICES STANDARD FLASHER TYPE (IMPORTED) HORIZONTAL RED & BLUE EMERGENCY LIGHT / LED FLASH LIGHT SIZE 1200 MM AND PA SYSTEM WITH SIREN/HOOTER SYSTEM AT LEAST 100 WATT OUTPUT. ALL EXTERNAL AND INTERNAL LIGHTING SHOULD BE LED’S | 1 |  |  |  |
| 15 | EXTERIOR ART DESIGN REFLECTING TAPES EMERGENCY ARTWORK ON ALL SIDES AND WORD ‘AMBULANCE’ ON FRONT, REAR AND BOTH SIDES, RED CRESCENT SIGN ON RIGHT AND LEFT WINDOWS AND “INSTITUTE OF SPACE TECHNOLOGY” MUST BE PLACED ON ALL THE SIDES. 2/3RD OF WINDOW GLASSES OF PATIENT COMPARTMENT TO BE COVERED WITH HIGH QUALITY COLORLESS, TRANSLUCENT IMPORTED FROSTED PAPER STICKER. | 1 |  |  |  |
| 16 | SPOILERS FOR FLASHERS SPECIALLY DESIGNED AERODYNAMIC MOLDED SPOILERS ON BOTH SIDES CONTAINING COMBINATION OF IMPORTED (NON-CHINESE) RED AND BLUE EEC FLASHERS WHICH ARE TUV CERTIFIED, AND JANPNESE FLOOD/SEARCH LIGHTS. | 1 |  |  |  |
| Medical Equipment for Ambulance: |
| S/No | Description | Qty | Yes | No | Alternative Offer |
| 1 | OXYGEN SUPPLY SYSTEM BRAND NEW IMPORTED OXYGEN CYLINDERS OF 10 LITER (WORKING PRESSURE: 150BAR, TEST PRESSURE: 250BAR, MATERIAL: 37MN) CAPACITY DULY TESTED, CERTIFIED AND FILLED BY GOVERNMENT APPROVED COMPANY WITH LOW OXYGEN ALARM, CHANGEOVER SYSTEM AND OXYGEN SUPPLY HOSE MADE OF RESILIENT, REINFORCED PLASTICS TO RETAIN IN ITS FLEXIBILITY IN COLD & UN-CRACKABLE IN HEAT UP TO 650C WITH BRASS SOCKETS AND CONNECTORS. THE PRESSURE REGULATOR (BRASS REGULATOR WITH BRITISH BULLNOSE AND 1 X 9/16” SELF-SEALING DISC OUTLETS (NO THERAPY FLOW) OF CYLINDERS WOULD BE OXYGEN WOULD BE DELIVERED THROUGH TWO OUTLETS ON THE PANEL WITH FLOW METERS AND FACE MASKS, ONE PRESSURE GAUGE SHOULD BE INSTALLED ON THE PANEL FOR MONITORING LINE PRESSURE | 1 |  |  |  |
| 2 | BP APPARATUS ACCOSON UK OR EQUIVALENT ONE BP APPARATUS WITH WALL MOUNTED LARGE DIAL TYPE 5/6” WITH STETHOSCOPE AND CUFFS | 1 |  |  |  |
| 3 | ONE FOLD STRETCHER TWO FOLD CE CERTIFIED WSX-D1 EVERISE (IMPORTED EURO STANDARD) BEDDING AREA MADE OF VERY STRONG WATER RESISTANCE PVC COATED WASHABLE ORANGE PARACHUTE CLOTH WITH ON IN IT AND SPECIAL CARRYING BAG MADE OF PARACHUTE AND STRUCTURE MADE OF 1 1/4” ROUND LIGHT-WEIGHT ALUMINUM ALLOY | 1 |  |  |  |
| 4 | MAIN AUTOMATIC LOADING STRETCHER CE CERTIFIED (IMPORTED EURO STANDARD) MAIN STRETCHER AUTOMATIC LOADING TYPE MADE OF ALUMINUM ALLOY/STAINLESS STEEL FRAME WITH 4+2 LOADING WHEELS AND ADJUSTABLE HEAD SECTION. THE MATTRESS SHOULD BE 2” THICK WITH BEST QUALITY FOAM, UPHOLSTERED WITH BEST QUALITY WASHABLE PVC COATED REXENE WITH PILLOW WITH TWO SAFETY BELTS AND TWO SIDE QUARDS. APPROX. SIZE 6’3” X 1’9” (LXW) WITH 3’ HT. IN STANDING POSITION AND 1’ HT. IN COLLAPSIBLE POSITION. LOCKING SYSTEM FOR STRETCHER AT FRONT AND RIGHT SIDE OF THE STRETCHER AND S.S SHEET ON THE FLOOR UNDER HIS STRETCHER. DETACHABLE IV ROD WITH THE STRETCHER | 1 |  |  |  |
| 5 | TRAUMA FIRST AID BAG ONE FIRST AID BAG MADE OF STRONG WATER RESISTANT PVC COATED WASHABLE RED PARACHUTE CLOTH WITH YELLOW REFLECTIVE STRIPS ON IT WITH APPROPRIATE PLACEMENT FOR AND INCLUDING THE FOLLOWING ITEMS:- WOODEN SPLINTS (2) CREPE BANDAGE (4”=016”=01) N/SALINE 0.9% (01) PYODINE SOLUTION (01) DRESSING SCISSOR (01) PYODINE POWDER (01) 4X4 GAUZE PAD (04) STICKING 1” (02) ANTISEPTIC WIPES 01 BOX | 1 |  |  |  |
| 6 | SUCTION APPARATUS MEDEL ITALY OR EQUIVALENT A RECHARGEABLE BATTERY OPERATED PORTABLE SUCTION DEVICE HAVING FREE OF MAINTENANCE VACCUM PUMPS, OUTCLAVABLE CONTAINERS EQUIPPED WITH AN ANTI-OVERFLOW VALVE AND VERY EFFECTIVE CLOSURE SYSTEM, INTERNALLY PROTECTED BY A TECHNO-POLYMER BODY EASY TO CLEAN, MAXIMUM VACUUM 600 MMHG, 23 L/MIN, BACTERIAL FILTER, SILICON & PVC TUBE, A SPECIAL HOLDING BRACKET AND SUCTION CATHETER | 1 |  |  |  |
| 7 | SPINE BOARD WITH HEAD IMMOBILIZER (IMPORTED EURO STANDARD) FLOATABLE SPINE BOARD THAT REDUCES THE POSSIBILITY OF CONTAMINATION DURING TRANSPORTATION OF THE PATIENT. HIGH DENSITY COMPLETELY RECYCLABLE POLYETHYLENE, MOLDED INTO A UNIQUE, ONE PIECE SHELL; THIS TECHNIQUE GUARANTEES THE B-BACK. THE EXCLUSIVE SPENCER CONTOUR HEAD IMMOBILIZER HAS BEEN STUDIED TO ASSURE THE OPTIMUM HEAD IMMOBILIZATION OF TRAUMATIZED PATIENTS | 1 |  |  |  |
| 8 | HARD CARRYING CASE ONE RESUSCITATION HARD CARRYING CASE MADE OF FG MATERIAL WITH WALL BRACKET. MADE OF STRONG ABS PLASTIC WHICH IS SHINY SMOOTH AND EASY TO CLEAN. A NEOPRENE SEAL IS FITTED TO MAKE THEM DUSTPROOF. A WALL BRACKET SECURE THEM ON ANY HOUSE WALL OR BOOTH. AMBO BAG HAVING SPECIAL VALVES WHICH SHOULD LIMIT THE EXCESSIVE FLOW OF GAS INTOTHE PATIENT AIRWAY SIGNIFICANTLY REDUCING THE RISK OF GASTRIC INSUFFLATIONS BY EFFECTIVELY LOWERING THE AIRWAY PRESSURE GENERATED. (OTWO CANADA/ SPENCER ITALY) ONE PORTABLE BP APPARATUS WITH STETHOSCOPE. HAND OPERATED MANUAL SUCTION APPARATUS. LARYNGOSCOPE WITH FOUR CURVED BLADES. PROFESSIONAL TORCH. MOUTH OPENER. ENDOTRACHEAL TUBES (DIFF) SIZE. AIRWAYS (DIFF) SIZES. GLOVES NICHIBAND. CONTAINMENT BAGS. | 1 |  |  |  |
| 9 | AED DEFIBRILLATOR (ZOLL USA/ DEFEBTCH, CANADA OR EQUIVALENT) AED DEFIBRILLATOR (AUTOMATIC EXTERNAL \DEFIBRILLATION WITH CPR FEEDBACK). THE AED MUST GIVE AUDIO AND VISUAL PROMPTS LIKE ‘PUSH HARDER’ AND ‘GOOD COMPRESSIONS’ TO GUIDE DURING CPR. (FDA/ CE CERTIFIED) THE AED SHOULD BE BIPHASIC WITH MAXIMUM ENERGY SETTING OF 200JOULES. THE AED SHOULD HAVE CONSUMER LITHIUM BATTERIES WITH FIVE YEAR SHELF-LIFE WHICH ARE READILY AVAILABLE AT LOCAL RETAIL OUTLETS AND MUST NOT BE DEFIBRILLATOR MANUFACTURER DEPENDENT. THE AED ELECTRODES SHOULD BE PREFERABLY REUSABLE. | 1 |  |  |  |
| 10 | PORTABLE VENTILATOR PROMPT (WEINMANN GERMANY/ALLIED UK OR EQUIVALENT) PORTABLE VENTILATOR CLEAR AND CONCISE VOICE PROMPTS LEAD TO SAFE VENTILATION WITH MAINTENANCE-FREE LITHIUM BATTERIES WITH A SERVICE LIFE OF AT LEAST TWO YEARS (NO CHARGING REQUIRED). THE OPTIMIZED OXYGEN USAGE WITH DEMAND-FLOW MODE. THE PATIENT RECEIVES OXYGEN DURING INSPIRATION; NO OXYGEN LOSS DURING EXHALATION AS CAUSED BY CONTINUOUS FLOW. THE CONTROLLED VENTILATION (IPPV) CAN BE SET BY SIMPLY TURNING THE ROTARY DIAL (TIME-CYCLED, VOLUME-CONSTANT). | 1 |  |  |  |
| 11 | ACCESSORIES OF VENTILATOR REUSABLE PATIENT HOSE SYSTEM PATIENT VALVE VENTILATION HOSE | 1 |  |  |  |
| 12 | SYRINGE CUTTER BEST QUALITY SYRINGE CUTTER | 1 |  |  |  |
| 13 | I.V HOOKS HOOKS FOR INTRAVENOUS INFUSION SET WILL BE PLACED | 1 |  |  |  |
| 14 | FIRE EXTINGUISHER 6 KG DRY CHEMICAL POWDER TYPE FIXED WITH BRACKET TO THE SIDE-WALL. 2 KG DCP IN DRIVER COMPARTMENT. | 1 |  |  |  |
| 15 | TRASH BOX STAINLESS STEEL WITH 6-8 LITERS CAPACITY, ATTACHED TO THE FLOOR AND WITH PLASTIC REMOVABLE AND WASHABLE INNER LINING | 1 |  |  |  |
| 16 | ANEROID SPHYGMOMANOMETER (UK/EURO STANDARD) TOP QUALITY WALL MOUNTED LARGE DIAL TYPE 6 WITH LATEST TECHNOLOGY AMBIDEX CUFF AND ONE STETHOSCOPE | 1 |  |  |  |
| 17 | NEBULIZER (BAURER GERMANY) ULTRASONIC NEBULIZER WITH RECHARGEABLE BATTERY. FOR THE TREATMENT OF THE UPPER AND LOWER RESPIRATORY TRACT. FOR USE IN COLDS, ASTHMA, RESPIRATORY DISEASE CAN BE DISINFECTED HIGH NEBULIZATION CAPACITY (>0.4 ML/MIN) – SHORT INHABITATION TIME HIGH PROPORTION OF RESPIRABLE PARTICLES STATE OF THE ART MEDICAL TECHNOLOGY: NEBULIZATION THROUGH HIGH-FREQUENT ULTRASONIC REPLACEMENT MEDICATION CONTAINER MAINS AND BATTERY OPERATION AUTOMATIC SWITCH-OFF FUNCTION AFTER 10 MIN ACCESSORIES: ADULT MASK, CHILDREN’S MASK, EXTENSION TUBE, RECHARGEABLE BATTERY, POWER ADAPTER, PROTECTIVE BAG | 1 |  |  |  |
| 18 | ECG MACHINE HIGH QUALITY PORTABLE BATTERY OPERATED ECG MACHINE WITH LCD AND PRINTER. | 1 |  |  |  |
| 19 | PATIENT MONITOR FOR VITAL SIGNS CE CERTIFIED EURO STANDARD MINIMUM 5.5 INCHES DISPLAY, LIGHT WEIGHT AND PORTABLE, SPO2 PULSE RATE AND MAIN STREAM CO2 MEASUREMENT WITH RECHARGEABLE BATTERY. | 1 |  |  |  |
| 20 | PORTABLE O2 CYLINDER WITH REGULATOR (IMPORTED EURO STANDARD) PORTABLE ALUMINUM OXYGEN CYLINDER WITH REGULATOR AND CARRYING BAG | 1 |  |  |  |
| 21 | SUCTION UNIT BATTERY OPERATED RECHARGEABLE BATTERY OPERATED SUCTION UNIT WITH MINIMUM 28 LITER/MIN AIRFLOW, 40 MINUTE BATTERY TIME, CLEAR BOTTLE AND ACCESSORIES | 1 |  |  |  |
| 22 | DELIVERY KIT GAMMA STERILIZED DELIVERY KIT CONTAINING FOLLOWING 9 ITEMS: 1X STERILIZED SURGICAL BLADE NO.10 (SINGLE EDGE) 2X STERILE CORD CLAMPS PLASTIC, INDIVIDUALLY PACKED 5X GAUZE PAD 12 PLY/LAYER 3" X 3" 10X COTTON SWAB. APPROX. 2 GM OR MORE. PACKED IN POLYTHENE BAG WITH LABEL OF MANUFACTURER 2X TOILET SOAP 50/60 GM 2X LATEX EXAMINATION GLOVES (PAIR) 1X TRANSPARENT PLASTIC SHEET 3FT X 3FT NOMINAL THICKNESS OF 0.2 MM OR MORE 2X POLYTHENE BAG 10 INCH X 10 INCH WITH NOMINAL THICKNESS OF 0.2 MM OR MORE 1X MUCUS EXTRACTOR | 1 |  |  |  |

Undertaking:

I/We, agree with the terms/conditions of tender inquiry. Items quoted are as per IST required specifications and rates are quoted accordingly. We have carefully read the terms and conditions mentioned in tender documents. We accept all these conditions. It is certified that the information furnished here in and as per the document submitted is true and correct and nothing has been concealed or tampered with. We have gone through all the conditions of tender and are liable to any punitive action for furnishing false information / documents.

Dated this _____ day of ____________________ 2018

Signature:

Name:

Designation:

Stamp:

(To be Printed On Company’s Letterhead)

Annexure – B

Financial Offer

LIT/IST Div/1718/3000000104

Date: 07-Jun-2018

Due On: 22-Jun-2018

| Installation Work (Fabrication of Ambulance Vehicle): |
| --- |
| S. No | Description/Specifications | QTY | Unit Price (to be quoted excluding GST) | Total Value |
| 1 | FLOOR FLOOR WOULD NOT BE PUNCTURED FOR INSTALLATION. FLOOR MADE OF WATER AND TERMITE PROOF HALF INCH THICK PLY-WOOD. THE FLOOR WOULD BE CARPETED WITH VINYL SHEET 2-3 MM THICK. THE FLOORING WOULD COVER ALL SIDES AND GAPS OF THE CABINETRY MAKING IT COMPLETELY WATER TIGHT. SPECIAL HARD POINTS OF 5MM THICKNESS MS SHEET MUST BE CREATED FOR INSTALLATION OF CHAIR, CABINET AND BENCH ETC. A SPECIAL S/S 16SWG SHEET WILL BE PROVIDED ON THE FLOOR UNDER THE STRETCHER AREA. | 1 |  |  |
| 2 | CEILING MADE OF FG MATERIAL WITH INTERNAL LED LIGHTS, ONE LONG GRIP HANDLE ALONG THE LENGTH OF CEILING MADE OF POWDER COATING PROPERLY REINFORCED AND OPTION FOR HOLDING TWO IV HOOKS WITH ANTI-SWING BELTS IN CEILING OR SIDE PANEL | 1 |  |  |
| 3 | PARTITION WALL PARTITION WALL BETWEEN THE DRIVER AND PATIENT COMPARTMENT WOULD BE MADE OF FG/ABS SHEET FIXED WITH STEEL BRACKETS TO SIDE WALLS. A SMALL WINDOW WOULD BE GIVEN IN BETWEEN THE DRIVER AND PATIENT COMPARTMENT WITH A BUILT-IN PLACE WITH GLASS FOR COMMUNICATION | 1 |  |  |
| 4 | SIDE UTILITY VERTICAL PANEL MADE OF ABS/FG MOLDED MATERIAL ON THE RIGHT SIDE DEDICATED O2 SYSTEM HAVING TWO OUTLETS, O2 ALARM SYSTEM, O2 CYLINDER CHANGEOVER SYSTEM AND ONE PRESSURE GAUGE FOR LINE PRESSURE. ALL SWITCHES AND SOCKETS FOR 12 V & 220 V, HAVING 220 INVERTER AT THE BACK | 1 |  |  |
| 5 | SIDE CABINET FOR EQUIPMENT MADE OF COMPLETELY FG MATERIAL ON RIGHT SIDE OF THE VEHICLE STARTING FORM THE END OF MAIN CABINET TO THE FULL LENGTH OF RIGHT SIDE, ONE TOP SHELF TO ACCOMMODATE AED, MONITOR, EMERGENCY BAG, SEARCH LIGHTS, ALL HAVING SPECIAL BRACKETS; LOWER SHELF HAVING SUCTION UNIT, ONE WIRED SHELF FOR ACCESSORIES AND OTHER ITEMS | 1 |  |  |
| 6 | ELECTRICAL WIRING FIRE RETARDANT AND HIGH TEMPERATURE RESISTANT ALL ELECTRICAL WIRING OF AMBULATORY COMPARTMENT WOULD BE OF WIRING HARNESS SYSTEM WITH NO JOINT AND TAPING HAVING ALL CONNECTOR SYSTEM AND HAVING A SEPARATE FUSE BOX FOR EACH OPERATION. ALL WIRES WOULD BE COLOR COATED AND WOULD BE OF FIRE RETARDANT AND HIGH TEMPERATURE RESISTANT MATERIAL. THERE WOULD BE A SEPARATE MAINTENANCE-FREE BATTERY OF 40 AMPS FOR AMBULATORY COMPARTMENT WHICH IS TO BE CHARGED THROUGH THE VEHICLE GENERATOR. THERE WOULD BE A 220 VOLTS INLET AND OUTLET SOCKETS FOR OUTSIDE SOURCE VOLTS | 1 |  |  |
| 7 | THREE TO FOUR PERSON BENCH MADE OF 3” SPECIAL POLYURETHANE FOAM AS USED IN AUTOMOBILE SEATS WITH APPROXIMATE 16” WIDTH, MADE OVER A STORAGE BOX OPEN-ABLE IN TWO PARTS. HEAT SEALED SEAT COVER OF BEST QUALITY ARTIFICIAL LEATHER, SEAT BELTS AND CUSHIONED BACK. ONE SIDE BAR MADE OF 1” M.S POWDER COATED PIPE | 1 |  |  |
| 8 | ONE PERSON CREW SEAT ONE CREW SEAT FIXED TO THE BODY OF THE VEHICLE WITH WAIST SEAT BELT AT THE HEAD END OF PATIENT WITH BEST QUALITY ARTIFICIAL LEATHER / FABRIC | 1 |  |  |
| 9 | OXYGEN SECTION TO ACCOMMODATE TWO M.F SIZE O2 CYLINDERS, WOULD HAVE DOUBLE STEEL BRACKETS WITH RUBBER PADDING. AN INTERNAL LIGHT WITH A DOOR SWITCH. FLOOR OF CABINET MATTED WITH VINYL FLOORING | 1 |  |  |
| 10 | WASH BASIN AND DUST BIN INTEGRATED WASH BASIN WITH WATER TAP AND WATER RESERVOIR WITH ELECTRIC WATER PUMP AND ELECTRIC SWITCH FOR TAP. DUST BIN IN THE CABINET | 1 |  |  |
| 11 | DRAWER SECTION HAVING 3 TO 4 DRAWERS. ALL DRAWERS MADE OF S/S SHEET MATERIAL. ALL DRAWERS HAVING SPECIAL SELF CLOSING & SELF RETAINING CATCHER HANDLE LOCK SYSTEM WHICH IS NOT SELF OPEN-ABLE DURING EMERGENCY DRIVING | 1 |  |  |
| 12 | SPOT LAMP SPOT LAMP WITH FLEXIBLE NECK FOR PATIENT EXAMINATION | 1 |  |  |
| 13 | BACK FLOOD LIGHT BACK SEARCH LIGHT, ADEQUATE INTERNAL LIGHTING FOR HANDLING THE PATIENTS AT NIGHT | 1 |  |  |
| 14 | EXTERNAL AND INTERNAL LIGHT AND WARNING DEVICES STANDARD FLASHER TYPE (IMPORTED) HORIZONTAL RED & BLUE EMERGENCY LIGHT / LED FLASH LIGHT SIZE 1200 MM AND PA SYSTEM WITH SIREN/HOOTER SYSTEM AT LEAST 100 WATT OUTPUT. ALL EXTERNAL AND INTERNAL LIGHTING SHOULD BE LED’S | 1 |  |  |
| 15 | EXTERIOR ART DESIGN REFLECTING TAPES EMERGENCY ARTWORK ON ALL SIDES AND WORD ‘AMBULANCE’ ON FRONT, REAR AND BOTH SIDES, RED CRESCENT SIGN ON RIGHT AND LEFT WINDOWS AND “INSTITUTE OF SPACE TECHNOLOGY” MUST BE PLACED ON ALL THE SIDES. 2/3RD OF WINDOW GLASSES OF PATIENT COMPARTMENT TO BE COVERED WITH HIGH QUALITY COLORLESS, TRANSLUCENT IMPORTED FROSTED PAPER STICKER. | 1 |  |  |
| 16 | SPOILERS FOR FLASHERS SPECIALLY DESIGNED AERODYNAMIC MOLDED SPOILERS ON BOTH SIDES CONTAINING COMBINATION OF IMPORTED (NON-CHINESE) RED AND BLUE EEC FLASHERS WHICH ARE TUV CERTIFIED, AND JANPNESE FLOOD/SEARCH LIGHTS. | 1 |  |  |
| Total Value: |  |
| GST Amount: |  |
| Net Total Amount: |  |
| Medical Equipment for Ambulance: |
| S. No | Description/Specifications | QTY | Unit Price (to be quoted excluding GST) | Total Value |
| 1 | OXYGEN SUPPLY SYSTEM BRAND NEW IMPORTED OXYGEN CYLINDERS OF 10 LITER (WORKING PRESSURE: 150BAR, TEST PRESSURE: 250BAR, MATERIAL: 37MN) CAPACITY DULY TESTED, CERTIFIED AND FILLED BY GOVERNMENT APPROVED COMPANY WITH LOW OXYGEN ALARM, CHANGEOVER SYSTEM AND OXYGEN SUPPLY HOSE MADE OF RESILIENT, REINFORCED PLASTICS TO RETAIN IN ITS FLEXIBILITY IN COLD & UN-CRACKABLE IN HEAT UP TO 650C WITH BRASS SOCKETS AND CONNECTORS. THE PRESSURE REGULATOR (BRASS REGULATOR WITH BRITISH BULLNOSE AND 1 X 9/16” SELF-SEALING DISC OUTLETS (NO THERAPY FLOW) OF CYLINDERS WOULD BE OXYGEN WOULD BE DELIVERED THROUGH TWO OUTLETS ON THE PANEL WITH FLOW METERS AND FACE MASKS, ONE PRESSURE GAUGE SHOULD BE INSTALLED ON THE PANEL FOR MONITORING LINE PRESSURE | 1 |  |  |
| 2 | BP APPARATUS ACCOSON UK OR EQUIVALENT ONE BP APPARATUS WITH WALL MOUNTED LARGE DIAL TYPE 5/6” WITH STETHOSCOPE AND CUFFS | 1 |  |  |
| 3 | ONE FOLD STRETCHER TWO FOLD CE CERTIFIED WSX-D1 EVERISE (IMPORTED EURO STANDARD) BEDDING AREA MADE OF VERY STRONG WATER RESISTANCE PVC COATED WASHABLE ORANGE PARACHUTE CLOTH WITH ON IN IT AND SPECIAL CARRYING BAG MADE OF PARACHUTE AND STRUCTURE MADE OF 1 1/4” ROUND LIGHT-WEIGHT ALUMINUM ALLOY | 1 |  |  |
| 4 | MAIN AUTOMATIC LOADING STRETCHER CE CERTIFIED (IMPORTED EURO STANDARD) MAIN STRETCHER AUTOMATIC LOADING TYPE MADE OF ALUMINUM ALLOY/STAINLESS STEEL FRAME WITH 4+2 LOADING WHEELS AND ADJUSTABLE HEAD SECTION. THE MATTRESS SHOULD BE 2” THICK WITH BEST QUALITY FOAM, UPHOLSTERED WITH BEST QUALITY WASHABLE PVC COATED REXENE WITH PILLOW WITH TWO SAFETY BELTS AND TWO SIDE QUARDS. APPROX. SIZE 6’3” X 1’9” (LXW) WITH 3’ HT. IN STANDING POSITION AND 1’ HT. IN COLLAPSIBLE POSITION. LOCKING SYSTEM FOR STRETCHER AT FRONT AND RIGHT SIDE OF THE STRETCHER AND S.S SHEET ON THE FLOOR UNDER HIS STRETCHER. DETACHABLE IV ROD WITH THE STRETCHER | 1 |  |  |
| 5 | TRAUMA FIRST AID BAG ONE FIRST AID BAG MADE OF STRONG WATER RESISTANT PVC COATED WASHABLE RED PARACHUTE CLOTH WITH YELLOW REFLECTIVE STRIPS ON IT WITH APPROPRIATE PLACEMENT FOR AND INCLUDING THE FOLLOWING ITEMS:- WOODEN SPLINTS (2) CREPE BANDAGE (4”=016”=01) N/SALINE 0.9% (01) PYODINE SOLUTION (01) DRESSING SCISSOR (01) PYODINE POWDER (01) 4X4 GAUZE PAD (04) STICKING 1” (02) ANTISEPTIC WIPES 01 BOX | 1 |  |  |
| 6 | SUCTION APPARATUS MEDEL ITALY OR EQUIVALENT A RECHARGEABLE BATTERY OPERATED PORTABLE SUCTION DEVICE HAVING FREE OF MAINTENANCE VACCUM PUMPS, OUTCLAVABLE CONTAINERS EQUIPPED WITH AN ANTI-OVERFLOW VALVE AND VERY EFFECTIVE CLOSURE SYSTEM, INTERNALLY PROTECTED BY A TECHNO-POLYMER BODY EASY TO CLEAN, MAXIMUM VACUUM 600 MMHG, 23 L/MIN, BACTERIAL FILTER, SILICON & PVC TUBE, A SPECIAL HOLDING BRACKET AND SUCTION CATHETER | 1 |  |  |
| 7 | SPINE BOARD WITH HEAD IMMOBILIZER (IMPORTED EURO STANDARD) FLOATABLE SPINE BOARD THAT REDUCES THE POSSIBILITY OF CONTAMINATION DURING TRANSPORTATION OF THE PATIENT. HIGH DENSITY COMPLETELY RECYCLABLE POLYETHYLENE, MOLDED INTO A UNIQUE, ONE PIECE SHELL; THIS TECHNIQUE GUARANTEES THE B-BACK. THE EXCLUSIVE SPENCER CONTOUR HEAD IMMOBILIZER HAS BEEN STUDIED TO ASSURE THE OPTIMUM HEAD IMMOBILIZATION OF TRAUMATIZED PATIENTS | 1 |  |  |
| 8 | HARD CARRYING CASE ONE RESUSCITATION HARD CARRYING CASE MADE OF FG MATERIAL WITH WALL BRACKET. MADE OF STRONG ABS PLASTIC WHICH IS SHINY SMOOTH AND EASY TO CLEAN. A NEOPRENE SEAL IS FITTED TO MAKE THEM DUSTPROOF. A WALL BRACKET SECURE THEM ON ANY HOUSE WALL OR BOOTH. AMBO BAG HAVING SPECIAL VALVES WHICH SHOULD LIMIT THE EXCESSIVE FLOW OF GAS INTOTHE PATIENT AIRWAY SIGNIFICANTLY REDUCING THE RISK OF GASTRIC INSUFFLATIONS BY EFFECTIVELY LOWERING THE AIRWAY PRESSURE GENERATED. (OTWO CANADA/ SPENCER ITALY) ONE PORTABLE BP APPARATUS WITH STETHOSCOPE. HAND OPERATED MANUAL SUCTION APPARATUS. LARYNGOSCOPE WITH FOUR CURVED BLADES. PROFESSIONAL TORCH. MOUTH OPENER. ENDOTRACHEAL TUBES (DIFF) SIZE. AIRWAYS (DIFF) SIZES. GLOVES NICHIBAND. CONTAINMENT BAGS. | 1 |  |  |
| 9 | AED DEFIBRILLATOR (ZOLL USA/ DEFEBTCH, CANADA OR EQUIVALENT) AED DEFIBRILLATOR (AUTOMATIC EXTERNAL \DEFIBRILLATION WITH CPR FEEDBACK). THE AED MUST GIVE AUDIO AND VISUAL PROMPTS LIKE ‘PUSH HARDER’ AND ‘GOOD COMPRESSIONS’ TO GUIDE DURING CPR. (FDA/ CE CERTIFIED) THE AED SHOULD BE BIPHASIC WITH MAXIMUM ENERGY SETTING OF 200JOULES. THE AED SHOULD HAVE CONSUMER LITHIUM BATTERIES WITH FIVE YEAR SHELF-LIFE WHICH ARE READILY AVAILABLE AT LOCAL RETAIL OUTLETS AND MUST NOT BE DEFIBRILLATOR MANUFACTURER DEPENDENT. THE AED ELECTRODES SHOULD BE PREFERABLY REUSABLE. | 1 |  |  |
| 10 | PORTABLE VENTILATOR PROMPT (WEINMANN GERMANY/ALLIED UK OR EQUIVALENT) PORTABLE VENTILATOR CLEAR AND CONCISE VOICE PROMPTS LEAD TO SAFE VENTILATION WITH MAINTENANCE-FREE LITHIUM BATTERIES WITH A SERVICE LIFE OF AT LEAST TWO YEARS (NO CHARGING REQUIRED). THE OPTIMIZED OXYGEN USAGE WITH DEMAND-FLOW MODE. THE PATIENT RECEIVES OXYGEN DURING INSPIRATION; NO OXYGEN LOSS DURING EXHALATION AS CAUSED BY CONTINUOUS FLOW. THE CONTROLLED VENTILATION (IPPV) CAN BE SET BY SIMPLY TURNING THE ROTARY DIAL (TIME-CYCLED, VOLUME-CONSTANT). | 1 |  |  |
| 11 | ACCESSORIES OF VENTILATOR REUSABLE PATIENT HOSE SYSTEM PATIENT VALVE VENTILATION HOSE | 1 |  |  |
| 12 | SYRINGE CUTTER BEST QUALITY SYRINGE CUTTER | 1 |  |  |
| 13 | I.V HOOKS HOOKS FOR INTRAVENOUS INFUSION SET WILL BE PLACED | 1 |  |  |
| 14 | FIRE EXTINGUISHER 6 KG DRY CHEMICAL POWDER TYPE FIXED WITH BRACKET TO THE SIDE-WALL. 2 KG DCP IN DRIVER COMPARTMENT. | 1 |  |  |
| 15 | TRASH BOX STAINLESS STEEL WITH 6-8 LITERS CAPACITY, ATTACHED TO THE FLOOR AND WITH PLASTIC REMOVABLE AND WASHABLE INNER LINING | 1 |  |  |
| 16 | ANEROID SPHYGMOMANOMETER (UK/EURO STANDARD) TOP QUALITY WALL MOUNTED LARGE DIAL TYPE 6 WITH LATEST TECHNOLOGY AMBIDEX CUFF AND ONE STETHOSCOPE | 1 |  |  |
| 17 | NEBULIZER (BAURER GERMANY) ULTRASONIC NEBULIZER WITH RECHARGEABLE BATTERY. FOR THE TREATMENT OF THE UPPER AND LOWER RESPIRATORY TRACT. FOR USE IN COLDS, ASTHMA, RESPIRATORY DISEASE CAN BE DISINFECTED HIGH NEBULIZATION CAPACITY (>0.4 ML/MIN) – SHORT INHABITATION TIME HIGH PROPORTION OF RESPIRABLE PARTICLES STATE OF THE ART MEDICAL TECHNOLOGY: NEBULIZATION THROUGH HIGH-FREQUENT ULTRASONIC REPLACEMENT MEDICATION CONTAINER MAINS AND BATTERY OPERATION AUTOMATIC SWITCH-OFF FUNCTION AFTER 10 MIN ACCESSORIES: ADULT MASK, CHILDREN’S MASK, EXTENSION TUBE, RECHARGEABLE BATTERY, POWER ADAPTER, PROTECTIVE BAG | 1 |  |  |
| 18 | ECG MACHINE HIGH QUALITY PORTABLE BATTERY OPERATED ECG MACHINE WITH LCD AND PRINTER. | 1 |  |  |
| 19 | PATIENT MONITOR FOR VITAL SIGNS CE CERTIFIED EURO STANDARD MINIMUM 5.5 INCHES DISPLAY, LIGHT WEIGHT AND PORTABLE, SPO2 PULSE RATE AND MAIN STREAM CO2 MEASUREMENT WITH RECHARGEABLE BATTERY. | 1 |  |  |
| 20 | PORTABLE O2 CYLINDER WITH REGULATOR (IMPORTED EURO STANDARD) PORTABLE ALUMINUM OXYGEN CYLINDER WITH REGULATOR AND CARRYING BAG | 1 |  |  |
| 21 | SUCTION UNIT BATTERY OPERATED RECHARGEABLE BATTERY OPERATED SUCTION UNIT WITH MINIMUM 28 LITER/MIN AIRFLOW, 40 MINUTE BATTERY TIME, CLEAR BOTTLE AND ACCESSORIES | 1 |  |  |
| 22 | DELIVERY KIT GAMMA STERILIZED DELIVERY KIT CONTAINING FOLLOWING 9 ITEMS: 1X STERILIZED SURGICAL BLADE NO.10 (SINGLE EDGE) 2X STERILE CORD CLAMPS PLASTIC, INDIVIDUALLY PACKED 5X GAUZE PAD 12 PLY/LAYER 3" X 3" 10X COTTON SWAB. APPROX. 2 GM OR MORE. PACKED IN POLYTHENE BAG WITH LABEL OF MANUFACTURER 2X TOILET SOAP 50/60 GM 2X LATEX EXAMINATION GLOVES (PAIR) 1X TRANSPARENT PLASTIC SHEET 3FT X 3FT NOMINAL THICKNESS OF 0.2 MM OR MORE 2X POLYTHENE BAG 10 INCH X 10 INCH WITH NOMINAL THICKNESS OF 0.2 MM OR MORE 1X MUCUS EXTRACTOR | 1 |  |  |
| Total Value: |  |
| GST Amount: |  |
| Net Total Amount: |  |

Undertaking:

I/We, agree with the terms/conditions of tender inquiry. Items quoted are as per IST required specifications and rates are quoted accordingly. We have carefully read the terms and conditions mentioned in tender documents. We accept all these conditions. It is certified that the information furnished here in and as per the document submitted is true and correct and nothing has been concealed or tampered with. We have gone through all the conditions of tender and are liable to any punitive action for furnishing false information / documents.

Dated this _____ day of ____________________ 2018

Signature:

Name:

Designation:

Stamp:

(To be Printed On Company’s Letterhead)

(To be Printed and typed On Company’s Letterhead)

(Annex-C)

LIT/IST Div/1718/3000000104

Date: 07-Jun-2018

Due On: 22-Jun-2018

Format for Bid

- Total Quoted Price 			:

- GST Amount (If applicable)		:

- Bid Bond Value				:

- Bid Bond PO/DD No.			:

- Bid Bond PO/DD dated			:

- Bid Bond PO/DD issued by

(Bank Name)				:

- Complete Technical specifications

& catalog attached				:			        YES  NO

- Single Offer Quoted

(Options are not allowed)			:			         YES  NO

- Delivery Schedule				:

- Validity of Acceptance			:

- Country of Origin				:

- Offered items are genuine and new	:			        YES  NO

- Performance Bond to be provided	:		                    YES  NO

- Installation & Commissioning offered	:			         YES  NO

- Guarantee/ Warranty			:	     		         YES  NO

- Guarantee/ Warranty contact details	:

E-mail					:

Ptcl #					:

Cell #					:

- Company profile/ capability for

completion of such projects attached	:			         YES  NO

- Company full address			:

Contact Person				:

E-mail					:

Cell #					:

- Acceptance of terms & conditions

of tender without any condition		:			         YES  NO

_____________

Sign & Seal

Annexure – D

Proposal Submission Form

LIT/IST Div/1718/3000000104

Date: 07-Jun-2018

Due On: 22-Jun-2018

Head of Procurement

Institute of Space Technology

Near CDA Rawat Toll Plaza, Islamabad Highway, Islamabad.

Subject: FABRICATION OF AMBULANCE (TOYOTA HIACE) FOR INSTITUTE OF SPACE TECHNOLOGY (IST) ISLAMABAD

Dear Sir,

A. Having examined the tender document and Appendixes we, the undersigned, in conformity with the said document, offer to provide the said items on terms of reference to be signed upon the award of contract for the sum indicated as per Price Schedule.

B. We undertake, if our proposal is accepted, to provide the items/services comprise in the contract within time frame specified, starting from the date of receipt of notification of award from the client Department / Office.

C. We agree to abide by this proposal for the period of 90 days (as per requirement of the project) from the date of bid opening and it shall remain binding upon us and may be accepted at any time before the expiration of that period.

D. We agree to execute a contract in the form to be communicated by the IST incorporating all agreements with such alterations or additions thereto as may be necessary to adapt such agreement to the circumstances of the standard.

E. We understand that you are not bound to accept a lowest or any bid you may receive, not to give any reason for rejection of any bid and that you will not defray any expenses incurred by us in biding.

Authorized Signatures with Official Seal